## Sheet1
filename | sheetname | sonde_id | field_visit_performed_by | parameter | field_date | post_deployment_date | exchange_time | post_deployment_time | post_cal_performed_by | verification_pre_clean | verification_post_clean | sonde_pre_clean | sonde_post_clean | sonde_post_deployment_standard_1 | standard_post_deployment_standard_1 | field_exchange_datetime | post_deployment_datetime | post_cal_time_elapsed_h | group | sonde_post_deployment_standard_2 | standard_post_deployment_standard_2 | field_check_verification_sonde | field_check_removed_sonde | field_check_new_installed_sonde | field_check_verifcation_date_time | post_calibration_staff | field_staff_second | fouling_error | fouling_error_pct | calibration_error | calibration_error_pct | total_error | total_error_pct | wagner_rating | year | month | flag_cal_pct | flag_cal
20220120TRPJORDRIFT | Primary | JOR | TB | SPC | 01/20/2022 | 01/20/2022 | 00:00 | 00:00 | TB | 393.1 | 392.9 | 393.4 | 388.1 | 1000 | 965 | 2022-01-20 08:00 | 01/20/2022 00:00 | 0 | MWQP | [8 empty] | 5.1 | 1.2974 | -35 | -3.5 | 40.1 | 4.7974 | Good | 2022 | 1 | out | out
DRIFT_TRP_2022_06_14 | Primary | JOR | TB | SPC | 06/14/2022 | 06/15/2022 | 00:00 | 00:00 | TB | 425 | 425 | 420 | 423 | 1000 | 977 | 2022-06-14 07:00 | 06/15/2022 00:00 | 24 | MWQP | [8 empty] | -3 | -0.70588 | -23 | -2.3 | 26 | 3.0059 | Good | 2022 | 6 | in | out
DRIFT_TRP_2022_07_06 v2 | Primary | PIP | JDC/CG | SPC | 07/06/2022 | 07/07/2022 | 00:00 | 00:00 | JDC | 538 | 538 | 536 | 541 | 1000 | 990 | 2022-07-06 07:00 | 07/07/2022 00:00 | 24 | MWQP | [8 empty] | -5 | -0.92937 | -10 | -1 | 15 | 1.9294 | Excellent | 2022 | 7 | in | out
DRIFT_TRP_2022_09_08 v2 | Primary | PIP | AC/CG | SPC | 09/08/2022 | 09/09/2022 | 00:00 | 00:00 | CG/AC | 449 | 449 | 451 | 450 | 1000 | 999 | 2022-09-08 07:00 | 09/09/2022 00:00 | 24 | MWQP | [8 empty] | 1 | 0.22272 | -1 | -0.1 | 2 | 0.32272 | Excellent | 2022 | 9 | in | in
DRIFT_TRP_2022_10_14 | Primary | JOR | TB/CG | SPC | 10/14/2022 | 10/17/2022 | 00:00 | 00:00 | CG | 510 | 510 | 509 | 508 | 1000 | 991 | 2022-10-14 07:00 | 10/17/2022 00:00 | 72 | MWQP | [8 empty] | 1 | 0.19608 | -9 | -0.9 | 10 | 1.0961 | Excellent | 2022 | 10 | in | out
DRIFT_TRP_2022_12_20 | Primary | RIC | AC/CG | SPC | 12/20/2022 |  | 00:00 | 00:00 | 44916 | 643 | 644 | 633 | 617 | 1000 | 975 | 2022-12-20 08:00 |  |  | MWQP | [8 empty] | 17 | 2.6439 | -25 | -2.5 | 42 | 5.1439 | Good | 2022 | 12 | in | out
DRIFT_TRP_8 | Primary | WES | TB | SPC | 08/10/2022 |  | 00:00 | 00:00 | 44783 | 443 | 443 | 443 | 448 | 1000 | 1002 | 2022-08-10 07:00 |  |  | MWQP | [8 empty] | -5 | -1.1287 | 2 | 0.2 | 7 | 1.3287 | Excellent | 2022 | 8 | in | in
HRO_Driftsheet_12_20_2022 | Primary | Jordan | CG/AC | SPC | 12/20/2022 | 12/21/2022 | 00:00 | 00:00 | AC | 686 | 687 | 682.5 | 682.3 | 1000 | 987 | 2022-12-20 08:00 | 12/21/2022 00:00 | 24 | MWQP | [8 empty] | 1.2 | 0.17493 | -13 | -1.3 | 14.2 | 1.4749 | Excellent | 2022 | 12 | in | out
HRO_DriftSheet_1_26_2022 | Primary | WES | JDC | SPC | 01/26/2022 | 01/26/2022 | 00:00 | 00:00 | JDC | 414 | 416 | 417 | 418 | 1000 | 1005.8 | 2022-01-26 08:00 | 01/26/2022 00:00 | 0 | MWQP | [8 empty] | 1 | 0.24155 | 5.8 | 0.58 | 6.8 | 0.82155 | Excellent | 2022 | 1 | in | out
HRO_Driftsheet_3_09_2022 | Primary | MAG | JDC | SPC | 03/09/2022 | 03/09/2022 | 00:00 | 00:00 | JDC | 462 | 463 | 460 | 464 | 1000 | 1007.1 | 2022-03-09 08:00 | 03/09/2022 00:00 | 0 | MWQP | [8 empty] | -3 | -0.64935 | 7.1 | 0.71 | 10.1 | 1.3594 | Excellent | 2022 | 3 | in | out
HRO_Driftsheet_4_08_2022 | Primary | BIR | JDC | SPC | 04/08/2022 | 04/08/2022 | 00:00 | 00:00 | JDC | 489 | 490 | 486 | 486 | 1000 | 995.3 | 2022-04-08 07:00 | 04/08/2022 00:00 | 0 | MWQP | [8 empty] | 1 | 0.2045 | -4.7 | -0.47 | 5.7 | 0.6745 | Excellent | 2022 | 4 | in | out
HRO_Driftsheet_5_19_2022 | Primary | WES | JDC | SPC | 05/19/2022 | 05/19/2022 | 00:00 | 00:00 | JDC | 438 | 438 | 443 | 436 | 1000 | 1013 | 2022-05-19 07:00 | 05/19/2022 00:00 | 0 | MWQP | [8 empty] | 7 | 1.5982 | 13 | 1.3 | 20 | 2.8982 | Excellent | 2022 | 5 | in | out
HRO_Driftsheet_7_6_2022 | Primary | RAM | JDC | SPC | 07/06/2022 | 07/07/2022 | 00:00 | 00:00 | JDC | 430 | 483 | 476 | 472 | 1000 | 989 | 2022-07-06 07:00 | 07/07/2022 00:00 | 24 | MWQP | [8 empty] | 57 | 13.2558 | -11 | -1.1 | 68 | 14.3558 | Fair | 2022 | 7 | in | out
HRO_Driftsheet_8_10_2022 | Primary | JOR | JDC | SPC | 08/10/2022 | 08/11/2022 | 00:00 | 00:00 | JDC | 501 | 501 | 503 | 487 | 1000 | 1004 | 2022-08-10 07:00 | 08/11/2022 00:00 | 24 | MWQP | [8 empty] | 16 | 3.1936 | 4 | 0.4 | 20 | 3.5936 | Good | 2022 | 8 | in | out
Oroville Sonde Drift calculations_20220113 | WQ-1 | OR-10 | SM/KM | SPC | 01/12/2022 | 01/13/2022 | 08:48 | 12:35 | SM | 114.4 | 114.5 | 113.05 | 113.01 | 1000 | 987.3 | 2022-01-12 16:48 | 01/13/2022 12:35 | 27.8 | NRO | [8 empty] | 0.14 | 0.12238 | -12.7 | -1.27 | 12.84 | 1.3924 | Excellent | 2022 | 1 | in | out
Oroville Sonde Drift calculations_20220224 | WQ-1 | OR-11 | PJ, SM | SPC | 02/23/2022 | 02/24/2022 | 10:00 | 10:45 | SM | 110.5 | 110.5 | 109.68 | 109.69 | 1000 | 991.8 | 2022-02-23 18:00 | 02/24/2022 10:45 | 24.8 | NRO | [8 empty] | -0.01 | -0.0090498 | -8.2 | -0.82 | 8.21 | 0.82905 | Excellent | 2022 | 2 | in | out
Oroville Sonde Drift calculations_20221130 | WQ-1 | OR-4 | SM/CB | SPC | 11/30/2022 | 12/03/2022 | 09:36 | 13:00 | SM | 115.9 | 116 | 115.63 | 117.14 | 1000 | 1001 | 2022-11-30 17:36 | 12/03/2022 13:00 | 75.4 | NRO | [8 empty] | -1.41 | -1.2166 | 1 | 0.1 | 2.41 | 1.3166 | Excellent | 2022 | 11 | in | in
Oroville Sonde Drift calculations_20221206 | WQ-4 | OR-4 | SM | SPC | 12/06/2022 | 12/06/2022 | 09:20 | 12:50 | SM | 107.8 | 107.7 | 108.3 | 108.7 | 1000 | 997.8 | 2022-12-06 17:20 | 12/06/2022 12:50 | 3.5 | NRO | [8 empty] | -0.5 | -0.46382 | -2.2 | -0.22 | 2.7 | 0.68382 | Excellent | 2022 | 12 | in | in
Thompson Meadow Sonde Drift Error calculations_20220512 | TC-Upper | TC-02M | PJ | SPC | 05/11/2022 | 05/12/2022 | 09:00 | 13:55 | PJ | 366 | 365.2 | 361.2 | 360.1 | 1000 | 995.7 | 2022-05-11 16:00 | 05/12/2022 13:55 | 28.9 | NRO | [8 empty] | 0.3 | 0.081967 | -4.3 | -0.43 | 4.6 | 0.51197 | Excellent | 2022 | 5 | in | out
Thompson Meadow Sonde Drift Error calculations_20220705 | TC-Upper | TC-1 | PJ | SPC | 06/30/2022 | 07/05/2022 | 07:38 | 09:05 | PJ | 652 | 652 | 660.7 | 662.8 | 1000 | 1009 | 2022-06-30 14:38 | 07/05/2022 09:05 | 121.5 | NRO | [8 empty] | -2.1 | -0.32209 | 9 | 0.9 | 11.1 | 1.2221 | Excellent | 2022 | 6 | in | out
Thompson Meadow Sonde Drift Error calculations_20220810 | TC-Upper | TC-02M | PJ | SPC | 08/10/2022 | 08/10/2022 | 07:10 | 17:11 | PJ | 761 | 761 | 756 | 755.6 | 1000 | 978.8 | 2022-08-10 14:10 | 08/10/2022 17:11 | 10 | NRO | [8 empty] | 0.4 | 0.052562 | -21.2 | -2.12 | 21.6 | 2.1726 | Excellent | 2022 | 8 | in | out
Thompson Meadow Sonde Drift Error calculations_20221023 | TC-Upper | TC-1 | PJ | SPC | 10/20/2022 | 10/23/2022 | 09:25 | 08:30 | PJ | 626 | 626 | 624.3 | 624 | 1000 | 999.4 | 2022-10-20 16:25 | 10/23/2022 08:30 | 71.1 | NRO | [8 empty] | 0.3 | 0.047923 | -0.6 | -0.06 | 0.9 | 0.10792 | Excellent | 2022 | 10 | in | in
| MHO20211201 | XO2-02 | BJones | SPC | [5 empty] | 555 | 555 | 554 | 556 | 2752.4 | 2767 | 2022-01-04 17:00 | 01/05/2022 16:15 | 31.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.36036 | 14.6 | 0.53045 | 16.6 | 0.89081 | Excellent | 2022 | 1 | in | out
| MRU20211201 | XO2-03 | BJones | SPC | [5 empty] | 618 | 618 | 618 | 618 | 2756 | 2767 | 2022-01-04 18:00 | 01/05/2022 06:00 | 20 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 11 | 0.39913 | 11 | 0.39913 | Excellent | 2022 | 1 | in | out
| TWA20211201 | XO2-48 | BJones | SPC | [5 empty] | 621 | 618 | 614 | 617 | 2743.7 | 2767 | 2022-01-04 18:45 | 01/05/2022 16:25 | 29.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6 | -0.96618 | 23.3 | 0.84922 | 29.3 | 1.8154 | Excellent | 2022 | 1 | in | out
| MRX20211201 | XO2-55 | BJones | SPC | [5 empty] | 861 | 859 | 856 | 859 | 2757.1 | 2767 | 2022-01-04 20:00 | 01/05/2022 15:40 | 27.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5 | -0.58072 | 9.9 | 0.35907 | 14.9 | 0.93979 | Excellent | 2022 | 1 | in | out
| ORM20211216 | XO2-10 | AMaguire | SPC | [5 empty] | 703 | 703 | 705 | 701 | 2760 | 2767 | 2022-01-05 19:00 | 01/06/2022 09:00 | 22 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4 | 0.56899 | 7 | 0.25362 | 11 | 0.82261 | Excellent | 2022 | 1 | in | out
| GLC20211214 | XO2-43 | AMaguire | SPC | [5 empty] | 703 | 702 | 704 | 704 | 2762 | 2767 | 2022-01-05 21:00 | 01/06/2022 09:00 | 20 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.14225 | 5 | 0.18103 | 6 | 0.32328 | Excellent | 2022 | 1 | in | out
| TSL20211117 | XO2-18 | TSalman | SPC | [5 empty] | 234.2 | 234.3 | 235.2 | 235.4 | 2755.2 | 2767 | 2022-01-06 17:45 | 01/07/2022 10:30 | 24.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.042699 | 11.8 | 0.42828 | 11.9 | 0.47098 | Excellent | 2022 | 1 | in | out
| FCT20211117 | XO2-60 | TSalman | SPC | [5 empty] | 235 | 235.1 | 240.9 | 240.5 | 2813 | 2767 | 2022-01-06 18:30 | 01/07/2022 09:07 | 22.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.21277 | -46 | -1.6353 | 46.5 | 1.848 | Excellent | 2022 | 1 | in | out
| FAL20211117 | XO2-66 | TSalman | SPC | [5 empty] | 235.9 | 236.1 | 237.2 | 237.7 | 2764 | 2767 | 2022-01-06 19:00 | 01/07/2022 09:40 | 22.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.12717 | 3 | 0.10854 | 3.3 | 0.23571 | Excellent | 2022 | 1 | in | out
| OSJ20211117 | XO2-16 | TSalman | SPC | [5 empty] | 236.4 | 236.5 | 236.9 | 236.6 | 2759 | 2767 | 2022-01-06 19:45 | 01/06/2022 16:57 | 5.2 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.1692 | 8 | 0.28996 | 8.4 | 0.45916 | Excellent | 2022 | 1 | in | out
| MOK20211117 | XO2-27 | TSalman | SPC | [5 empty] | 186.5 | 186.6 | 187.5 | 187.7 | 2760.7 | 2767 | 2022-01-06 20:15 | 01/07/2022 08:35 | 20.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.053619 | 6.3 | 0.2282 | 6.4 | 0.28182 | Excellent | 2022 | 1 | in | out
| ORI20211214 | XO2-35 | AMaguire | SPC | [5 empty] | 720 | 723 | 721 | 723 | 2763 | 2767 | 2022-01-11 18:30 | 01/12/2022 09:30 | 23 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.13889 | 4 | 0.14477 | 5 | 0.28366 | Excellent | 2022 | 1 | in | out
| WCI20211214 | XO2-40 | AMaguire | SPC | [5 empty] | 722 | 723 | 717 | 722 | 2763 | 2767 | 2022-01-11 19:00 | 01/12/2022 10:10 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.55402 | 4 | 0.14477 | 8 | 0.69879 | Excellent | 2022 | 1 | in | out
| VCU20211214 | XO2-64 | AMaguire | SPC | [5 empty] | 722 | 721 | 719 | 718 | 2758 | 2767 | 2022-01-11 19:45 | 01/12/2022 09:50 | 22.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 9 | 0.32632 | 9 | 0.32632 | Excellent | 2022 | 1 | in | out
| TRB820211208 | XO2-08 | LDavis | SPC | [5 empty] | 116.6 | 116.6 | 113 | 112.9 | 2750.7 | 2767 | 2022-01-12 16:45 | 01/13/2022 08:55 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.085763 | 16.3 | 0.59258 | 16.4 | 0.67834 | Excellent | 2022 | 1 | in | out
| SJD20211201 | XO2-04 | DColvin | SPC | [5 empty] | 660 | 660 | 658 | 658 | 2746 | 2767 | 2022-01-12 18:00 | 01/13/2022 15:00 | 29 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 21 | 0.76475 | 21 | 0.76475 | Excellent | 2022 | 1 | in | out
| OH120211201 | XO2-12 | DColvin | SPC | [5 empty] | 660 | 660 | 657 | 656 | 2744 | 2767 | 2022-01-12 18:15 | 01/12/2022 15:00 | 4.8 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.15152 | 23 | 0.83819 | 24 | 0.98971 | Excellent | 2022 | 1 | in | out
| TRTT20211208 | XO2-13 | LDavis | SPC | [5 empty] | 116 | 116.1 | 115.9 | 116 | 2760.5 | 2767 | 2022-01-12 18:15 | 01/13/2022 09:53 | 23.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 6.5 | 0.23546 | 6.5 | 0.23546 | Excellent | 2022 | 1 | in | out
| GLE20211201 | XO2-34 | DColvin | SPC | [5 empty] | 663 | 663 | 658 | 659 | 2748.3 | 2767 | 2022-01-12 19:15 | 01/12/2022 15:00 | 3.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.15083 | 18.7 | 0.68042 | 19.7 | 0.83125 | Excellent | 2022 | 1 | in | out
| PDC20211201 | XO2-14 | DColvin | SPC | [5 empty] | 664 | 665 | 665 | 663 | 2764.1 | 2767 | 2022-01-12 19:45 | 01/12/2022 15:00 | 3.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.45181 | 2.9 | 0.10492 | 5.9 | 0.55672 | Excellent | 2022 | 1 | in | in
| FAR20211208 | XO2-56 | LDavis | SPC | [5 empty] | 116.2 | 115.9 | 116.9 | 116.8 | 2763.3 | 2767 | 2022-01-12 19:45 | 01/13/2022 10:16 | 22.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.17212 | 3.7 | 0.1339 | 3.9 | 0.30601 | Excellent | 2022 | 1 | in | out
| SGA20211201 | XO2-28 | DColvin | SPC | [5 empty] | 681 | 670 | 675 | 668 | 2753.1 | 2767 | 2022-01-12 20:00 | 01/13/2022 15:00 | 27 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.58737 | 13.9 | 0.50489 | 17.9 | 1.0923 | Excellent | 2022 | 1 | in | out
| SXS20211130 | XO2-52 | PScott | SPC | [5 empty] | 192.1 | 192.3 | 192.8 | 192.5 | 2753 | 2767 | 2022-01-13 18:00 | 01/13/2022 14:20 | 4.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.26028 | 14 | 0.50854 | 14.5 | 0.76882 | Excellent | 2022 | 1 | in | out
| SOI20211130 | XO2-59 | PScott | SPC | [5 empty] | 192.5 | 192.8 | 193 | 193.2 | 2756.8 | 2767 | 2022-01-13 18:30 | 01/13/2022 14:20 | 3.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.051948 | 10.2 | 0.36999 | 10.3 | 0.42194 | Excellent | 2022 | 1 | in | out
| MIR20211130 | XO2-61 | PScott | SPC | [5 empty] | 189.3 | 189.1 | 188.3 | 188.3 | 2755.6 | 2767 | 2022-01-13 19:00 | 01/13/2022 14:20 | 3.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.10565 | 11.4 | 0.4137 | 11.6 | 0.51936 | Excellent | 2022 | 1 | in | out
| TRN20211208 | XO2-20 | TSalman | SPC | [5 empty] | 358.9 | 358.9 | 357.9 | 358.6 | 2775 | 2767 | 2022-01-19 18:30 | 01/29/2022 11:40 | 241.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.7 | -0.19504 | -8 | -0.28829 | 8.7 | 0.48333 | Excellent | 2022 | 1 | in | out
| MRU20220104 | XO2-42 | BJones | SPC | [5 empty] | 228.2 | 227 | 234.2 | 227.1 | 2770 | 2767 | 2022-01-19 19:06 | 01/19/2022 14:40 | 3.6 | NCRO | FALSE | FALSE | TRUE |  | TRUE |  |  |  | 5.9 | 2.5855 | -3 | -0.1083 | 8.9 | 2.6938 | Excellent | 2022 | 1 | in | out
| HLT20211208 | XO2-33 | TSalman | SPC | [5 empty] | 358.6 | 358.8 | 357.6 | 358 | 2767 | 2767 | 2022-01-19 19:30 | 01/20/2022 11:55 | 24.4 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.055772 | 0 | 0 | 0.2 | 0.055772 | Excellent | 2022 | 1 | in | in
| HOL20211208 | XO2-58 | TSalman | SPC | [5 empty] | 313.6 | 313.8 | 314.9 | 314.3 | 2772.2 | 2767 | 2022-01-19 20:45 | 01/20/2022 12:05 | 23.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.2551 | -5.2 | -0.18758 | 6 | 0.44268 | Excellent | 2022 | 1 | in | out
| OBI20211208 | XO2-69 | TSalman | SPC | [5 empty] | 314 | 314.1 | 312.8 | 314 | 2761.5 | 2767 | 2022-01-19 21:15 | 01/20/2022 13:15 | 24 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.1 | -0.35032 | 5.5 | 0.19917 | 6.6 | 0.54949 | Excellent | 2022 | 1 | in | out
| OSJ20220106 | XO2-29 | AMaguire | SPC | [5 empty] | 222.6 | 222.5 | 229.6 | 226.1 | 2769 | 2767 | 2022-01-21 16:45 | 01/21/2022 08:48 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.4 | 1.5274 | -2 | -0.072228 | 5.4 | 1.5996 | Excellent | 2022 | 1 | in | in
| SGA20220112 | XO2-50 | DColvin | SPC | [5 empty] | 1460 | 1460 | 1455 | 1461 | 2765 | 2767 | 2022-01-25 18:30 | 01/26/2022 12:00 | 25.5 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6 | -0.41096 | 2 | 0.072333 | 8 | 0.48329 | Excellent | 2022 | 1 | in | in
| LAB20211208 | XO2-15 | LDavis | SPC | [5 empty] | 80.7 | 80.5 | 80.2 | 80.9 | 2762.6 | 2767 | 2022-01-25 18:45 | 01/26/2022 10:06 | 23.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.9 | -1.1152 | 4.4 | 0.15927 | 5.3 | 1.2745 | Excellent | 2022 | 1 | in | out
| PDC20220112 | XO2-41 | DColvin | SPC | [5 empty] | 1456 | 1455 | 1456 | 1455 | 2772.4 | 2767 | 2022-01-25 19:15 | 01/26/2022 12:15 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -5.4 | -0.19478 | 5.4 | 0.19478 | Excellent | 2022 | 1 | in | out
| WDC20211209 | XO2-23 | CWeber | SPC | [5 empty] | 361.1 | 355 | 358.8 | 353 | 2763 | 2767 | 2022-01-26 17:15 | 01/28/2022 14:00 | 52.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.083079 | 4 | 0.14477 | 4.3 | 0.22785 | Excellent | 2022 | 1 | in | out
| RSCC20211209 | XO2-47 | CWeber | SPC | [5 empty] | 356 | 356.1 | 355.7 | 359.3 | 2765.5 | 2767 | 2022-01-26 18:00 | 01/28/2022 14:00 | 52 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3.5 | -0.98315 | 1.5 | 0.05424 | 5 | 1.0374 | Excellent | 2022 | 1 | in | in
| DRB20211209 | XO2-31 | CWeber | SPC | [5 empty] | 357.1 | 356.5 | 352.1 | 352.3 | 2739.5 | 2767 | 2022-01-26 18:45 | 01/28/2022 14:00 | 51.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.22403 | 27.5 | 1.0038 | 28.3 | 1.2279 | Excellent | 2022 | 1 | in | out
| HCHM20211209 | XO2-01 |  | SPC | [5 empty] | 356.4 | 356.8 | 355 | 356.5 | 2762.2 | 2767 | 2022-01-26 19:15 | 01/28/2022 14:00 | 50.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.1 | -0.30864 | 4.8 | 0.17377 | 5.9 | 0.48242 | Excellent | 2022 | 1 | in | out
| ORM20220105 | XO2-05 | AMaguire | SPC | [5 empty] | 971 | 971 | 975 | 977 | 2765 | 2767 | 2022-02-02 18:15 | 02/03/2022 09:30 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.20597 | 2 | 0.072333 | 4 | 0.27831 | Excellent | 2022 | 2 | in | in
| ODM20211102 | XO2-49 | AMaguire | SPC | [5 empty] | 968 | 969 | 949 | 976 | 2766 | 2767 | 2022-02-02 18:45 | 02/03/2022 09:30 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -26 | -2.686 | 1 | 0.036153 | 27 | 2.7221 | Excellent | 2022 | 2 | in | in
| GLC20220105 | XO2-63 | AMaguire | SPC | [5 empty] | 970 | 969 | 975 | 977 | 2772 | 2767 | 2022-02-02 19:30 | 02/03/2022 09:30 | 22 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.30928 | -5 | -0.18038 | 8 | 0.48965 | Excellent | 2022 | 2 | in | out
| BET20211230 | XO2-22 | JFrantzich | SPC | [5 empty] | 337.5 | 336.5 | 343.6 | 339.4 | 2752.2 | 2767 | 2022-02-03 20:00 | 02/03/2022 16:00 | 4 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.2 | 0.94815 | 14.8 | 0.53775 | 18 | 1.4859 | Excellent | 2022 | 2 | in | out
| BLP20211230 | XO2-09 | JFrantzich | SPC | [5 empty] | 419.1 | 419.1 | 420.2 | 421.2 | 2766.7 | 2767 | 2022-02-03 21:15 | 02/03/2022 16:20 | 3.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.23861 | 0.3 | 0.010843 | 1.3 | 0.24945 | Excellent | 2022 | 2 | in | in
| MHO20220104 | XO2-32 | BJones | SPC | [5 empty] | 918 | 917 | 900 | 920 | 2756.4 | 2767 | 2022-02-08 16:45 | 02/09/2022 13:40 | 28.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -21 | -2.2876 | 10.6 | 0.38456 | 31.6 | 2.6721 | Excellent | 2022 | 2 | in | out
| ORI20220111 | XO2-10 | AMaguire | SPC | [5 empty] | 389.7 | 390.1 | 392.9 | 393.5 | 2775 | 2767 | 2022-02-08 18:00 | 02/08/2022 16:30 | 6.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.051322 | -8 | -0.28829 | 8.2 | 0.33961 | Excellent | 2022 | 2 | in | out
| MRU20220119 | XO2-34 | BJones | SPC | [5 empty] | 836 | 835 | 838 | 839 | 2755.4 | 2767 | 2022-02-08 18:15 | 02/09/2022 11:45 | 25.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.23923 | 11.6 | 0.42099 | 13.6 | 0.66023 | Excellent | 2022 | 2 | in | out
| WCI20220111 | XO2-44 | AMaguire | SPC | [5 empty] | 390.3 | 390.2 | 392.4 | 393.1 | 2773.6 | 2767 | 2022-02-08 18:30 | 02/09/2022 09:20 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.20497 | -6.6 | -0.23796 | 7.4 | 0.44293 | Excellent | 2022 | 2 | in | out
| TWA20220104 | XO2-07 | BJones | SPC | [5 empty] | 840 | 836 | 843 | 842 | 2762.1 | 2767 | 2022-02-08 19:00 | 02/08/2022 11:03 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.35714 | 4.9 | 0.1774 | 7.9 | 0.53454 | Excellent | 2022 | 2 | in | out
| VCU20220111 | XO2-43 | AMaguire | SPC | [5 empty] | 391 | 390.6 | 392.7 | 392.8 | 2766.7 | 2767 | 2022-02-08 19:15 | 02/09/2022 09:00 | 21.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.12788 | 0.3 | 0.010843 | 0.8 | 0.13872 | Excellent | 2022 | 2 | in | in
| MRX20220104 | XO2-06 | BJones | SPC | [5 empty] | 394.9 | 394.2 | 396.6 | 397.8 | 2752.1 | 2767 | 2022-02-08 20:30 | 02/09/2022 11:30 | 23 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.9 | -0.48113 | 14.9 | 0.5414 | 16.8 | 1.0225 | Excellent | 2022 | 2 | in | out
| TRB820220112 | XO2-18 | LDavis | SPC | [5 empty] | 113.7 | 113.8 | 114.8 | 115.2 | 2772.9 | 2767 | 2022-02-09 16:30 | 02/10/2022 09:25 | 24.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.26385 | -5.9 | -0.21277 | 6.2 | 0.47663 | Excellent | 2022 | 2 | in | out
| TRTT20220112 | XO2-46 | LDavis | SPC | [5 empty] | 114.1 | 114.2 | 114.9 | 115 | 2768.2 | 2767 | 2022-02-09 18:00 | 02/10/2022 09:38 | 23.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.4211e-14 | 1.2455e-14 | -1.2 | -0.043349 | 1.2 | 0.043349 | Excellent | 2022 | 2 | in | in
| FAR20220112 | XO2-60 | LDavis | SPC | [5 empty] | 114.1 | 113.9 | 115.7 | 115 | 2770.3 | 2767 | 2022-02-09 19:30 | 02/10/2022 09:55 | 22.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.43821 | -3.3 | -0.11912 | 3.8 | 0.55733 | Excellent | 2022 | 2 | in | out
| OSJ20220121 | XO2-33 | TSalman | SPC | [5 empty] | 268.6 | 268.6 | 270.8 | 270.2 | 2762 | 2767 | 2022-02-10 17:45 | 02/10/2022 09:49 | 0.1 | NCRO | FALSE | FALSE | TRUE |  | TRUE |  |  |  | 0.6 | 0.22338 | 5 | 0.18103 | 5.6 | 0.40441 | Excellent | 2022 | 2 | in | out
| LIS20220103 | ND-3 | EHuynh | SPC | [5 empty] | 714 | 716 | 713 | 713 | 2767.4 | 2767 | 2022-02-14 20:30 | 02/14/2022 15:30 | 3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.28011 | -0.4 | -0.014454 | 2.4 | 0.29457 | Excellent | 2022 | 2 | in | in
| TRN20220119 | XO2-35 | TSalman | SPC | [5 empty] | 620 | 621 | 621 | 621 | 2783.1 | 2767 | 2022-02-15 19:15 | 02/16/2022 10:30 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.16129 | -16.1 | -0.57849 | 17.1 | 0.73978 | Excellent | 2022 | 2 | in | out
| HLT20220119 | XO2-66 | TSalman | SPC | [5 empty] | 619 | 619 | 617 | 618 | 2787.5 | 2767 | 2022-02-15 20:00 | 02/16/2022 11:20 | 23.3 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16155 | -20.5 | -0.73543 | 21.5 | 0.89698 | Excellent | 2022 | 2 | in | out
| ORQ20220119 | XO2-59 | TSalman | SPC | [5 empty] | 620 | 620 | 616 | 620 | 2796.7 | 2767 | 2022-02-15 20:30 | 02/16/2022 13:30 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.64516 | -29.7 | -1.062 | 33.7 | 1.7071 | Excellent | 2022 | 2 | in | out
| OBI20220119 | XO2-70 | TSalman | SPC | [5 empty] | 619 | 619 | 623 | 621 | 2784.6 | 2767 | 2022-02-15 21:00 | 02/16/2022 14:00 | 25 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.3231 | -17.6 | -0.63205 | 19.6 | 0.95515 | Excellent | 2022 | 2 | in | out
| WDC20220125 | XO2-14 | CWeber | SPC | [5 empty] | 339.9 | 340 | 340.7 | 341.7 | 2770.3 | 2767 | 2022-02-16 16:15 | 02/17/2022 16:30 | 32.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.9 | -0.26478 | -3.3 | -0.11912 | 4.2 | 0.3839 | Excellent | 2022 | 2 | in | out
| RSCC20220125 | XO2-45 | CWeber | SPC | [5 empty] | 340.3 | 340.4 | 343.4 | 344.2 | 2788.3 | 2767 | 2022-02-16 16:45 | 02/17/2022 16:45 | 32 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.7 | -0.2057 | -21.3 | -0.76391 | 22 | 0.96961 | Excellent | 2022 | 2 | in | out
| DRB20220125 | XO2-21 | CWeber | SPC | [5 empty] | 340.9 | 341 | 341.3 | 341.5 | 2764.1 | 2767 | 2022-02-16 17:30 | 02/17/2022 16:00 | 30.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.029334 | 2.9 | 0.10492 | 3 | 0.13425 | Excellent | 2022 | 2 | in | in
| HCHM20220125 | XO2-04 | CWeber | SPC | [5 empty] | 341.2 | 341.4 | 340 | 341.9 | 2772.6 | 2767 | 2022-02-16 17:45 | 02/17/2022 09:00 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.7 | -0.49824 | -5.6 | -0.20198 | 7.3 | 0.70022 | Excellent | 2022 | 2 | in | out
| PDC20220125 | XO2-08 | DColvin | SPC | [5 empty] | 608 | 608 | 608 | 608 | 2779 | 2767 | 2022-02-16 20:45 | 02/17/2022 12:15 | 23.5 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -12 | -0.43181 | 12 | 0.43181 | Excellent | 2022 | 2 | in | out
| SXS20220113 | XO2-16 | PScott | SPC | [5 empty] | 179.6 | 179.7 | 180.1 | 180.2 | 2777 | 2767 | 2022-02-17 18:00 | 02/17/2022 14:39 | 4.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -10 | -0.3601 | 10 | 0.3601 | Excellent | 2022 | 2 | in | out
| SOI20220113 | XO2-27 | PScott | SPC | [5 empty] | 180.1 | 180.3 | 181.5 | 181.6 | 2776.6 | 2767 | 2022-02-17 18:30 | 02/17/2022 14:50 | 4.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.055525 | -9.6 | -0.34575 | 9.7 | 0.40127 | Excellent | 2022 | 2 | in | out
| MIR20220113 | XO2-64 | PScott | SPC | [5 empty] | 181 | 180.8 | 181.9 | 181.9 | 2777.3 | 2767 | 2022-02-17 19:00 | 02/17/2022 16:45 | 5.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.1105 | -10.3 | -0.37086 | 10.5 | 0.48136 | Excellent | 2022 | 2 | in | out
| TSL20220106 | XO2-26 | TSalman | SPC | [5 empty] | 247.5 | 247.5 | 249.3 | 249.3 | 2771.3 | 2767 | 2022-02-22 18:15 | 02/22/2022 10:19 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -4.3 | -0.15516 | 4.3 | 0.15516 | Excellent | 2022 | 2 | in | out
| FCT20220106 | XO2-53 | TSalman | SPC | [5 empty] | 248.1 | 248.1 | 249.6 | 249.4 | 2774.1 | 2767 | 2022-02-22 18:45 | 02/22/2022 10:00 | -0.8 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.080613 | -7.1 | -0.25594 | 7.3 | 0.33655 | Excellent | 2022 | 2 | in | out
| FAL20220106 | XO2-19 | TSalman | SPC | [5 empty] | 248.9 | 248.9 | 250.2 | 249.6 | 2776.9 | 2767 | 2022-02-22 19:30 | 02/22/2022 10:00 | -1.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.24106 | -9.9 | -0.35651 | 10.5 | 0.59757 | Excellent | 2022 | 2 | in | out
| MOK20220106 | XO2-54 | TSalman | SPC | [5 empty] | 248.9 | 248.8 | 250.8 | 250.9 | 2774.4 | 2767 | 2022-02-22 20:15 | 02/22/2022 10:00 | -2.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.080354 | -7.4 | -0.26672 | 7.6 | 0.34708 | Excellent | 2022 | 2 | in | out
| SJD20220112 | XO2-25 | DColvin | SPC | [5 empty] | 488.3 | 488.4 | 489.3 | 490.4 | 2770.1 | 2767 | 2022-03-01 19:00 | 03/01/2022 16:05 | 5.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.20479 | -3.1 | -0.11191 | 4.1 | 0.3167 | Excellent | 2022 | 3 | in | out
| OH120220112 | XO2-02 | DColvin | SPC | [5 empty] | 489 | 488.9 | 491.5 | 492.2 | 2778 | 2767 | 2022-03-01 19:30 | 03/01/2022 15:00 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.1636 | -11 | -0.39597 | 11.8 | 0.55957 | Excellent | 2022 | 3 | in | out
| ORX20220112 | XO2-03 | DColvin | SPC | [5 empty] | 480.5 | 482 | 481.9 | 484.5 | 2771 | 2767 | 2022-03-01 20:00 | 03/01/2022 16:35 | 4.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.1 | -0.22893 | -4 | -0.14435 | 5.1 | 0.37328 | Excellent | 2022 | 3 | in | out
| GLE20220112 | XO2-48 | DColvin | SPC | [5 empty] | 483.9 | 483.2 | 485.4 | 485.8 | 2768 | 2767 | 2022-03-01 20:30 | 03/01/2022 15:00 | 2.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.1 | -0.22732 | -1 | -0.036127 | 2.1 | 0.26345 | Excellent | 2022 | 3 | in | in
| GLC20220202 | XO2-61 | AMaguire | SPC | [5 empty] | 458.7 | 459.8 | 460.5 | 461.1 | 2761.4 | 2767 | 2022-03-08 18:00 | 03/09/2022 11:00 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.109 | 5.6 | 0.2028 | 6.1 | 0.3118 | Excellent | 2022 | 3 | in | out
| ORI20220208 | XO2-05 | AMaguire | SPC | [5 empty] | 460.4 | 460.5 | 461.8 | 460.3 | 2762.4 | 2767 | 2022-03-08 18:30 | 03/09/2022 11:30 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.6 | 0.34752 | 4.6 | 0.16652 | 6.2 | 0.51405 | Excellent | 2022 | 3 | in | out
| WCI20220208 | XO2-49 | AMaguire | SPC | [5 empty] | 460.9 | 460.9 | 460.3 | 461 | 2762.9 | 2767 | 2022-03-08 19:00 | 03/09/2022 10:45 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.7 | -0.15188 | 4.1 | 0.14839 | 4.8 | 0.30027 | Excellent | 2022 | 3 | in | out
| VCU20220208 | XO2-63 | AMaguire | SPC | [5 empty] | 461.6 | 461.7 | 460.3 | 460.9 | 2759 | 2767 | 2022-03-08 19:30 | 03/09/2022 12:45 | 25.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.10832 | 8 | 0.28996 | 8.5 | 0.39828 | Excellent | 2022 | 3 | in | out
| ORM20220202 | XO2-20 | AMaguire | SPC | [5 empty] | 594 | 593 | 593 | 595 | 2765.6 | 2767 | 2022-03-08 20:15 | 03/09/2022 13:15 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.50505 | 1.4 | 0.050622 | 4.4 | 0.55567 | Excellent | 2022 | 3 | in | in
| ODM20220202 | XO2-36 | AMaguire | SPC | [5 empty] | 592 | 593 | 584 | 595 | 2759.6 | 2767 | 2022-03-08 20:45 | 03/09/2022 12:15 | 23.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -10 | -1.6892 | 7.4 | 0.26815 | 17.4 | 1.9573 | Excellent | 2022 | 3 | in | out
| TRB820220209 | XO2-15 | LDavis | SPC | [5 empty] | 113.2 | 114.1 | 114.7 | 115.8 | 2765.7 | 2767 | 2022-03-09 16:30 | 03/10/2022 12:22 | 27.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.17668 | 1.3 | 0.047004 | 1.5 | 0.22368 | Excellent | 2022 | 3 | in | in
| TRTT20220209 | XO2-22 | LDavis | SPC | [5 empty] | 115.6 | 115.5 | 117.1 | 117.1 | 2782.6 | 2767 | 2022-03-09 17:30 | 03/10/2022 12:32 | 27 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.086505 | -15.6 | -0.56063 | 15.7 | 0.64713 | Excellent | 2022 | 3 | in | out
| HLT20220215 | XO2-37 | TSalman | SPC | [5 empty] | 326.5 | 326.6 | 327.8 | 328.5 | 2767.2 | 2767 | 2022-03-09 18:30 | 03/09/2022 16:00 | 5.5 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.6 | -0.18377 | -0.2 | -0.0072275 | 0.8 | 0.19099 | Excellent | 2022 | 3 | in | in
| FAR20220209 | XO2-69 | LDavis | SPC | [5 empty] | 115.4 | 115.5 | 117.3 | 117.1 | 2768.3 | 2767 | 2022-03-09 18:30 | 03/10/2022 12:07 | 25.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.25997 | -1.3 | -0.04696 | 1.6 | 0.30693 | Excellent | 2022 | 3 | in | in
| ORQ20220215 | XO2-11 | TSalman | SPC | [5 empty] | 275.1 | 275.1 | 279.5 | 276.8 | 2774.4 | 2767 | 2022-03-09 19:00 | 03/10/2022 15:15 | 28.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.7 | 0.98146 | -7.4 | -0.26672 | 10.1 | 1.2482 | Excellent | 2022 | 3 | in | out
| HOL20220203 | XO2-58 | TSalman | SPC | [5 empty] | 327.7 | 327.8 | 328 | 324.3 | 2763.8 | 2767 | 2022-03-09 20:00 | 03/09/2022 14:40 | 2.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.8 | 1.1596 | 3.2 | 0.11578 | 7 | 1.2754 | Excellent | 2022 | 3 | in | out
| OBI20220215 | XO2-65 | TSalman | SPC | [5 empty] | 327.8 | 328.6 | 329.9 | 329.4 | 2773.7 | 2767 | 2022-03-09 20:30 | 03/09/2022 16:00 | 3.5 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.3 | 0.39658 | -6.7 | -0.24155 | 8 | 0.63814 | Excellent | 2022 | 3 | in | out
| LIS20220214 | ND-1 | EHuynh | SPC | [5 empty] | 857 | 857 | 862 | 860 | 2781.8 | 2767 | 2022-03-14 22:00 | 03/15/2022 15:08 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.23337 | -14.8 | -0.53203 | 16.8 | 0.7654 | Excellent | 2022 | 3 | in | out
| MHO20220208 | XO2-55 | BJones | SPC | [5 empty] | 742 | 742 | 739 | 742 | 2773.5 | 2767 | 2022-03-15 15:15 | 03/16/2022 02:15 | 18 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.40431 | -6.5 | -0.23436 | 9.5 | 0.63867 | Excellent | 2022 | 3 | in | out
| MRU20220208 | XO2-42 | BJones | SPC | [5 empty] | 928 | 928 | 932 | 932 | 2766.7 | 2767 | 2022-03-15 16:15 | 03/16/2022 14:10 | 28.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 0.3 | 0.010843 | 0.3 | 0.010843 | Excellent | 2022 | 3 | in | in
| TWA20220208 | XO2-24 | BJones | SPC | [5 empty] | 929 | 928 | 923 | 930 | 2769.9 | 2767 | 2022-03-15 17:00 | 03/16/2022 14:35 | 28.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -8 | -0.86114 | -2.9 | -0.1047 | 10.9 | 0.96584 | Excellent | 2022 | 3 | in | in
| BET20220203 | XO2-13 | JFrantzich | SPC | [5 empty] | 310.3 | 310.3 | 311.2 | 312.2 | 2768.9 | 2767 | 2022-03-15 17:30 | 03/16/2022 14:50 | 28.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.32227 | -1.9 | -0.068619 | 2.9 | 0.39089 | Excellent | 2022 | 3 | in | in
| MRX20220208 | XO2-40 | BJones | SPC | [5 empty] | 496.3 | 496.1 | 497.1 | 493 | 2765.6 | 2767 | 2022-03-15 18:15 | 03/16/2022 14:25 | 27.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.9 | 0.78582 | 1.4 | 0.050622 | 5.3 | 0.83644 | Excellent | 2022 | 3 | in | in
| BLP20220203 | XO2-12 | JFrantzich | SPC | [5 empty] | 315.1 | 315.2 | 315.1 | 317.5 | 2777.7 | 2767 | 2022-03-15 20:00 | 03/15/2022 16:00 | 3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2.3 | -0.72993 | -10.7 | -0.38521 | 13 | 1.1151 | Excellent | 2022 | 3 | in | out
| PDC20220216 | XO2-47 | DColvin | SPC | [5 empty] | 805 | 805 | 804 | 805 | 2765.5 | 2767 | 2022-03-16 17:46 | 03/17/2022 11:00 | 24.2 | NCRO | FALSE | FALSE | TRUE |  | TRUE |  |  |  | -1 | -0.12422 | 1.5 | 0.05424 | 2.5 | 0.17846 | Excellent | 2022 | 3 | in | in
| SGA20220216 | XO2-09 | DColvin | SPC | [5 empty] | 807 | 806 | 806 | 800 | 2678 | 2767 | 2022-03-16 18:15 | 03/17/2022 11:15 | 24 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5 | 0.61958 | 89 | 3.3234 | 94 | 3.943 | Good | 2022 | 3 | out | out
| SXS20220217 | XO2-56 | PScott | SPC | [5 empty] | 166 | 166 | 167.4 | 167.6 | 2761 | 2767 | 2022-03-17 16:00 | 03/18/2022 15:00 | 30 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.12048 | 6 | 0.21731 | 6.2 | 0.33779 | Excellent | 2022 | 3 | in | out
| SOI20220217 | XO2-18 | PScott | SPC | [5 empty] | 166.1 | 166.2 | 167 | 166.7 | 2760 | 2767 | 2022-03-17 16:30 | 03/18/2022 15:00 | 29.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.24082 | 7 | 0.25362 | 7.4 | 0.49444 | Excellent | 2022 | 3 | in | out
| MIR20220217 | XO2-34 | PScott | SPC | [5 empty] | 152 | 152 | 152.3 | 152.2 | 2756.7 | 2767 | 2022-03-17 17:00 | 03/18/2022 15:00 | 29 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.065789 | 10.3 | 0.37364 | 10.4 | 0.43942 | Excellent | 2022 | 3 | in | out
| LAB20220125 | XO2-57 | LDavis | SPC | [5 empty] | 81.4 | 81.6 | 81.5 | 81.9 | 2764.1 | 2767 | 2022-03-22 15:15 | 03/23/2022 09:58 | 25.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.2457 | 2.9 | 0.10492 | 3.1 | 0.35062 | Excellent | 2022 | 3 | in | in
| WDC20220216 | XO2-23 | CWeber | SPC | [5 empty] | 338.1 | 338.9 | 344.8 | 339.7 | 2762.2 | 2767 | 2022-03-23 14:15 | 03/23/2022 14:45 | 7.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5.9 | 1.745 | 4.8 | 0.17377 | 10.7 | 1.9188 | Excellent | 2022 | 3 | in | out
| RSCC20220216 | XO2-50 | CWeber | SPC | [5 empty] | 339.9 | 339.2 | 339.2 | 339.7 | 2759.3 | 2767 | 2022-03-23 15:00 | 03/23/2022 14:45 | 6.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.2 | -0.35305 | 7.7 | 0.27906 | 8.9 | 0.6321 | Excellent | 2022 | 3 | in | out
| DRB20220216 | XO2-31 | CWeber | SPC | [5 empty] | 339.9 | 339.8 | 339.6 | 340.3 | 2759.6 | 2767 | 2022-03-23 15:30 | 03/23/2022 14:45 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.23536 | 7.4 | 0.26815 | 8.2 | 0.50352 | Excellent | 2022 | 3 | in | out
| HCHM20220216 | XO2-01 | CWeber | SPC | [5 empty] | 340.4 | 340.8 | 341.4 | 341.8 | 2761.3 | 2767 | 2022-03-23 16:00 | 03/23/2022 14:45 | 5.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 5.7 | 0.20642 | 5.7 | 0.20642 | Excellent | 2022 | 3 | in | out
| TSL20220222 | XO2-66 | TSalman | SPC | [5 empty] | 318.7 | 320.5 | 319.9 | 321.2 | 2756.1 | 2767 | 2022-03-24 16:30 | 03/25/2022 14:40 | 29.2 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.15689 | 10.9 | 0.39549 | 11.4 | 0.55237 | Excellent | 2022 | 3 | in | out
| FAL20220222 | XO2-35 | TSalman | SPC | [5 empty] | 321.3 | 321.4 | 321.6 | 322 | 2762.6 | 2767 | 2022-03-24 17:00 | 03/24/2022 16:00 | 6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.093371 | 4.4 | 0.15927 | 4.7 | 0.25264 | Excellent | 2022 | 3 | in | out
| FCT20220222 | XO2-70 | TSalman | SPC | [5 empty] | 322.2 | 321.9 | 322.5 | 322.7 | 2758.7 | 2767 | 2022-03-24 17:30 | 03/25/2022 14:50 | 28.3 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.15518 | 8.3 | 0.30087 | 8.8 | 0.45605 | Excellent | 2022 | 3 | in | out
| OSJ20220210 | XO2-43 | TSalman | SPC | [5 empty] | 322.4 | 322.3 | 322.7 | 325.3 | 2778.5 | 2767 | 2022-03-24 18:15 | 03/25/2022 15:15 | 28 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2.7 | -0.83747 | -11.5 | -0.41389 | 14.2 | 1.2514 | Excellent | 2022 | 3 | in | out
| MOK20220222 | XO2-59 | TSalman | SPC | [5 empty] | 321.7 | 322.6 | 319.5 | 320.5 | 2738.3 | 2767 | 2022-03-24 19:00 | 03/25/2022 14:20 | 26.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.031085 | 28.7 | 1.0481 | 28.8 | 1.0792 | Excellent | 2022 | 3 | in | out
| WCI20220308 | XO2-48 | AMaguire | SPC | [5 empty] | 441.1 | 441.1 | 438.6 | 439.3 | 2756 | 2767 | 2022-03-29 16:30 | 03/29/2022 13:00 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.7 | -0.15869 | 11 | 0.39913 | 11.7 | 0.55782 | Excellent | 2022 | 3 | in | out
| LIS20220322 | ND-4 | EHuynh | SPC | [5 empty] | 737 | 737 | 734 | 736 | 2759.5 | 2767 | 2022-04-04 17:45 | 04/04/2022 13:50 | 3.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.27137 | 7.5 | 0.27179 | 9.5 | 0.54316 | Excellent | 2022 | 4 | in | out
| GLC20220308 | XO2-19 | AMaguire | SPC | [5 empty] | 434.5 | 424.4 | 421.3 | 421.6 | 2759.4 | 2767 | 2022-04-05 16:30 | 04/05/2022 13:37 | 4.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -10.4 | -2.3936 | 7.6 | 0.27542 | 18 | 2.669 | Excellent | 2022 | 4 | in | out
| ODM20220308 | XO2-60 | AMaguire | SPC | [5 empty] | 425 | 425 | 422 | 418.3 | 2758 | 2767 | 2022-04-05 17:00 | 04/05/2022 14:00 | 4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.7 | 0.87059 | 9 | 0.32632 | 12.7 | 1.1969 | Excellent | 2022 | 4 | in | out
| ORM20220308 | XO2-64 | AMaguire | SPC | [5 empty] | 426.3 | 426.8 | 425.3 | 426.3 | 2769.3 | 2767 | 2022-04-05 17:30 | 04/05/2022 13:54 | 3.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.11729 | -2.3 | -0.083053 | 2.8 | 0.20034 | Excellent | 2022 | 4 | in | in
| GLE20220301 | XO2-32 | DColvin | SPC | [5 empty] | 769 | 769 | 768 | 770 | 2759.2 | 2767 | 2022-04-07 18:00 | 04/08/2022 11:45 | 24.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.26008 | 7.8 | 0.28269 | 9.8 | 0.54277 | Excellent | 2022 | 4 | in | out
| ORX20220301 | XO2-10 | DColvin | SPC | [5 empty] | 770 | 770 | 767 | 768 | 2761.9 | 2767 | 2022-04-07 18:30 | 04/08/2022 12:35 | 25.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.12987 | 5.1 | 0.18466 | 6.1 | 0.31453 | Excellent | 2022 | 4 | in | out
| OH120220301 | XO2-06 | DColvin | SPC | [5 empty] | 773 | 772 | 771 | 771 | 2754.8 | 2767 | 2022-04-07 19:00 | 04/07/2022 12:02 | 0 | NCRO | FALSE | FALSE | TRUE |  | TRUE |  |  |  | -1 | -0.12937 | 12.2 | 0.44286 | 13.2 | 0.57223 | Excellent | 2022 | 4 | in | out
| SJD20220301 | XO2-45 | DColvin | SPC | [5 empty] | 773 | 772 | 770 | 773 | 2761.4 | 2767 | 2022-04-07 19:15 | 04/08/2022 12:55 | 24.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.51746 | 5.6 | 0.2028 | 9.6 | 0.72026 | Excellent | 2022 | 4 | in | out
| WCI20220329 | XO2-24 | AMaguire | SPC | [5 empty] | 806 | 807 | 808 | 809 | 2771 | 2767 | 2022-04-12 16:30 | 04/12/2022 15:40 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -4 | -0.14435 | 4 | 0.14435 | Excellent | 2022 | 4 | in | out
| VCU20220308 | XO2-46 | AMaguire | SPC | [5 empty] | 805 | 806 | 797 | 799 | 2749 | 2767 | 2022-04-12 17:00 | 04/12/2022 15:00 | 5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.12422 | 18 | 0.65478 | 19 | 0.77901 | Excellent | 2022 | 4 | in | out
| ORI20220308 | XO2-27 | AMaguire | SPC | [5 empty] | 805 | 805 | 801 | 803 | 2761 | 2767 | 2022-04-12 17:30 | 04/12/2022 15:25 | 4.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.24845 | 6 | 0.21731 | 8 | 0.46576 | Excellent | 2022 | 4 | in | out
| TRB820220309 | XO2-16 | LDavis | SPC | [5 empty] | 95.2 | 95.3 | 95 | 95.2 | 2743.8 | 2767 | 2022-04-18 15:00 | 04/19/2022 08:50 | 24.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.10504 | 23.2 | 0.84554 | 23.3 | 0.95058 | Excellent | 2022 | 4 | in | out
| TRTT20220309 | XO2-26 | LDavis | SPC | [5 empty] | 104.5 | 104.7 | 104.7 | 104.3 | 2740.5 | 2767 | 2022-04-18 16:15 | 04/19/2022 09:01 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.57416 | 26.5 | 0.96698 | 27.1 | 1.5411 | Excellent | 2022 | 4 | in | out
| LAB20220322 | XO2-15 | LDavis | SPC | [5 empty] | 105.1 | 105.4 | 105.9 | 105.7 | 2753.7 | 2767 | 2022-04-18 17:00 | 04/19/2022 08:37 | 22.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.47574 | 13.3 | 0.48299 | 13.8 | 0.95872 | Excellent | 2022 | 4 | in | out
| FAR20220309 | XO2-53 | LDavis | SPC | [5 empty] | 106 | 106.2 | 106.5 | 106.4 | 2731 | 2767 | 2022-04-18 18:00 | 04/18/2022 14:37 | 3.6 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.28302 | 36 | 1.3182 | 36.3 | 1.6012 | Excellent | 2022 | 4 | in | out
| MHO20220315 | XO2-20 | BJones | SPC | [5 empty] | 695 | 696 | 686 | 688 | 2739.1 | 2767 | 2022-04-19 15:00 | 04/20/2022 10:20 | 26.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.14388 | 27.9 | 1.0186 | 28.9 | 1.1625 | Excellent | 2022 | 4 | in | out
| MRU20220315 | XO2-05 | BJones | SPC | [5 empty] | 1162 | 1162 | 1149 | 1145 | 2752.1 | 2767 | 2022-04-19 16:00 | 04/20/2022 09:20 | 24.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4 | 0.34423 | 14.9 | 0.5414 | 18.9 | 0.88564 | Excellent | 2022 | 4 | in | out
| TWA20220315 | XO2-36 | BJones | SPC | [5 empty] | 1163 | 1162 | 1150 | 1150 | 2755.1 | 2767 | 2022-04-19 16:45 | 04/20/2022 09:35 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.085985 | 11.9 | 0.43193 | 12.9 | 0.51791 | Excellent | 2022 | 4 | in | out
| TRN20220215 | XO2-44 | TSalman | SPC | [5 empty] | 532 | 533 | 528 | 530 | 2752 | 2767 | 2022-04-19 16:45 | 04/20/2022 13:57 | 28.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18797 | 15 | 0.54506 | 16 | 0.73303 | Excellent | 2022 | 4 | in | out
| BET20220315 | XO2-41 | JFrantzich | SPC | [5 empty] | 312.5 | 312.8 | 312.4 | 312.9 | 2755.9 | 2767 | 2022-04-19 17:15 | 04/20/2022 10:18 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.064 | 11.1 | 0.40277 | 11.3 | 0.46677 | Excellent | 2022 | 4 | in | out
| HLT20220309 | XO2-03 | TSalman | SPC | [5 empty] | 533 | 533 | 527 | 529 | 2757.4 | 2767 | 2022-04-19 17:45 | 04/20/2022 14:30 | 27.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.37523 | 9.6 | 0.34815 | 11.6 | 0.72339 | Excellent | 2022 | 4 | in | out
| MRX20220315 | XO2-49 | BJones | SPC | [5 empty] | 413 | 413 | 409.1 | 411.4 | 2757.5 | 2767 | 2022-04-19 18:15 | 04/20/2022 10:00 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2.3 | -0.5569 | 9.5 | 0.34451 | 11.8 | 0.90142 | Excellent | 2022 | 4 | in | out
| ORQ20220309 | XO2-54 | TSalman | SPC | [5 empty] | 530 | 532 | 528 | 524 | 2748.6 | 2767 | 2022-04-19 18:30 | 04/20/2022 16:00 | 28.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 6 | 1.1321 | 18.4 | 0.66943 | 24.4 | 1.8015 | Excellent | 2022 | 4 | in | out
| BLP20220315 | XO2-61 | JFrantzich | SPC | [5 empty] | 573 | 574 | 569 | 574 | 2754 | 2767 | 2022-04-19 18:30 | 04/19/2022 16:17 | 4.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.69808 | 13 | 0.47204 | 17 | 1.1701 | Excellent | 2022 | 4 | in | out
| HOL20220309 | XO2-25 | TSalman | SPC | [5 empty] | 532 | 532 | 527 | 524 | 2752.6 | 2767 | 2022-04-19 19:00 | 04/20/2022 15:00 | 27 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.56391 | 14.4 | 0.52314 | 17.4 | 1.0871 | Excellent | 2022 | 4 | in | out
| OBI20220309 | XO2-21 | TSalman | SPC | [5 empty] | 530 | 531 | 528 | 528 | 2750 | 2767 | 2022-04-19 19:45 | 04/20/2022 16:57 | 28.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.18868 | 17 | 0.61818 | 18 | 0.80686 | Excellent | 2022 | 4 | in | out
| SGA20220316 | XO2-72 | DColvin | SPC | [5 empty] | 702 | 702 | 695 | 698 | 2707 | 2767 | 2022-04-20 16:45 | 04/20/2022 00:00 | -9.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.42735 | 60 | 2.2165 | 63 | 2.6438 | Excellent | 2022 | 4 | in | out
| PDC20220316 | XO2-73 | DColvin | SPC | [5 empty] | 703 | 703 | 699 | 699 | 2740 | 2767 | 2022-04-20 17:15 | 04/20/2022 16:40 | 6.4 | NCRO | FALSE | FALSE | TRUE | TRUE |  |  |  |  | 0 | 0 | 27 | 0.9854 | 27 | 0.9854 | Excellent | 2022 | 4 | in | out
| GLE20220407 | XO2-19 | DColvin | SPC | [5 empty] | 704 | 704 | 704 | 705 | 2750 | 2767 | 2022-04-20 17:45 | 04/21/2022 16:20 | 29.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.14205 | 17 | 0.61818 | 18 | 0.76023 | Excellent | 2022 | 4 | in | out
| WDC20220323 | XO2-08 | CWeber | SPC | [5 empty] | 319.7 | 321 | 318.4 | 319.1 | 2749.9 | 2767 | 2022-04-25 15:15 | 04/26/2022 16:00 | 31.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.18768 | 17.1 | 0.62184 | 17.7 | 0.80952 | Excellent | 2022 | 4 | in | out
| RSCC20220323 | XO2-30 | CWeber | SPC | [5 empty] | 321.9 | 322.4 | 320.7 | 320.7 | 2748.2 | 2767 | 2022-04-25 16:00 | 04/26/2022 16:00 | 31 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.15533 | 18.8 | 0.68408 | 19.3 | 0.83941 | Excellent | 2022 | 4 | in | out
| DRB20220323 | XO2-14 | CWeber | SPC | [5 empty] | 323.4 | 323.5 | 316 | 321.4 | 2757.3 | 2767 | 2022-04-25 16:45 | 04/25/2022 16:00 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5.3 | -1.6388 | 9.7 | 0.35179 | 15 | 1.9906 | Excellent | 2022 | 4 | in | out
| HCHM20220323 | XO2-04 | CWeber | SPC | [5 empty] | 323 | 323 | 320.9 | 320 | 2746 | 2767 | 2022-04-25 17:15 | 04/26/2022 16:00 | 29.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.9 | 0.27864 | 21 | 0.76475 | 21.9 | 1.0434 | Excellent | 2022 | 4 | in | out
| MOK20220324 | XO2-18 | TSalman | SPC | [5 empty] | 148.8 | 148.9 | 148.8 | 148.8 | 2749.2 | 2767 | 2022-04-26 16:15 | 04/27/2022 09:17 | 24 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.067204 | 17.8 | 0.64746 | 17.9 | 0.71467 | Excellent | 2022 | 4 | in | out
| OSJ20220324 | XO2-55 | TSalman | SPC | [5 empty] | 149.8 | 150.1 | 151.6 | 149.4 | 2755.9 | 2767 | 2022-04-26 16:45 | 04/27/2022 12:30 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.5 | 1.6689 | 11.1 | 0.40277 | 13.6 | 2.0717 | Excellent | 2022 | 4 | in | out
| FAL20220324 | XO2-37 | TSalman | SPC | [5 empty] | 152.1 | 152.3 | 152.1 | 152.1 | 2754 | 2767 | 2022-04-26 17:30 | 04/27/2022 15:40 | 29.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.13149 | 13 | 0.47204 | 13.2 | 0.60353 | Excellent | 2022 | 4 | in | out
| FCT20220324 | XO2-56 | TSalman | SPC | [5 empty] | 153.3 | 153.4 | 155 | 154.1 | 2751 | 2767 | 2022-04-26 18:00 | 04/27/2022 11:10 | 24.2 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.65232 | 16 | 0.58161 | 17 | 1.2339 | Excellent | 2022 | 4 | in | out
| TSL20220324 | XO2-13 | TSalman | SPC | [5 empty] | 155.5 | 154.9 | 156.1 | 151.2 | 2742.8 | 2767 | 2022-04-26 18:45 | 04/27/2022 15:45 | 28 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4.3 | 2.7653 | 24.2 | 0.88231 | 28.5 | 3.6476 | Good | 2022 | 4 | in | out
| MIR20220317 | XO2-02 | PScott | SPC | [5 empty] | 169 | 169.1 | 168.3 | 168.3 | 2748 | 2767 | 2022-04-28 16:15 | 04/29/2022 10:00 | 24.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.059172 | 19 | 0.69141 | 19.1 | 0.75058 | Excellent | 2022 | 4 | in | out
| SOI20220317 | XO2-69 | PScott | SPC | [5 empty] | 169.5 | 169.8 | 169.8 | 169.5 | 2748.4 | 2767 | 2022-04-28 16:45 | 04/29/2022 10:00 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.35398 | 18.6 | 0.67676 | 19.2 | 1.0307 | Excellent | 2022 | 4 | in | out
| SXS20220317 | XO2-11 | PScott | SPC | [5 empty] | 169.9 | 170.1 | 169.1 | 169.2 | 2745.3 | 2767 | 2022-04-28 17:15 | 04/29/2022 10:00 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.058858 | 21.7 | 0.79044 | 21.8 | 0.8493 | Excellent | 2022 | 4 | in | out
| LIS20220404 | ND-1 | EHuynh | SPC | [5 empty] | 605 | 606 | 607 | 605 | 2767.1 | 2767 | 2022-05-02 17:30 | 05/03/2022 10:50 | 24.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.49587 | -0.1 | -0.0036139 | 3.1 | 0.49948 | Excellent | 2022 | 5 | in | in
| SJD20220407 | XO2-47 | BJones | SPC | [5 empty] | 271.8 | 272.1 | 270.9 | 271.1 | 2766 | 2767 | 2022-05-03 15:00 | 05/04/2022 12:25 | 28.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.036792 | 1 | 0.036153 | 1.1 | 0.072945 | Excellent | 2022 | 5 | in | in
| OH120220407 | XO2-64 | BJones | SPC | [5 empty] | 272.1 | 272.4 | 273 | 271.8 | 2775 | 2767 | 2022-05-03 15:30 | 05/04/2022 12:15 | 27.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.5 | 0.55127 | -8 | -0.28829 | 9.5 | 0.83956 | Excellent | 2022 | 5 | in | out
| ORX20220407 | XO2-60 | BJones | SPC | [5 empty] | 272.9 | 273.1 | 273 | 273 | 2770 | 2767 | 2022-05-03 16:00 | 05/04/2022 11:55 | 26.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.073287 | -3 | -0.1083 | 3.2 | 0.18159 | Excellent | 2022 | 5 | in | out
| GLE20220420 | XO2-09 | BJones | SPC | [5 empty] | 274.3 | 274 | 274.5 | 274.5 | 2777.8 | 2767 | 2022-05-03 16:30 | 05/03/2022 09:38 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.10937 | -10.8 | -0.3888 | 11.1 | 0.49817 | Excellent | 2022 | 5 | in | out
| GLC20220405 | XO2-42 | AMaguire | SPC | [5 empty] | 396.8 | 397 | 397 | 396.9 | 2774 | 2767 | 2022-05-04 15:15 | 05/04/2022 13:34 | 5.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.075605 | -7 | -0.25234 | 7.3 | 0.32795 | Excellent | 2022 | 5 | in | out
| ORM20220405 | XO2-48 | AMaguire | SPC | [5 empty] | 399.9 | 399.9 | 401.4 | 400.8 | 2777.3 | 2767 | 2022-05-04 16:45 | 05/04/2022 13:00 | 3.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.15004 | -10.3 | -0.37086 | 10.9 | 0.5209 | Excellent | 2022 | 5 | in | out
| ODM20220405 | XO2-12 | AMaguire | SPC | [5 empty] | 400.9 | 400.9 | 399.7 | 398.1 | 2762 | 2767 | 2022-05-04 17:15 | 05/04/2022 13:00 | 2.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.6 | 0.3991 | 5 | 0.18103 | 6.6 | 0.58013 | Excellent | 2022 | 5 | in | out
| ORI20220412 | XO2-06 | AMaguire | SPC | [5 empty] | 410.9 | 411.2 | 411.1 | 413.2 | 2769.3 | 2767 | 2022-05-10 16:45 | 05/11/2022 08:25 | 22.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.8 | -0.43806 | -2.3 | -0.083053 | 4.1 | 0.52112 | Excellent | 2022 | 5 | in | in
| WCI20220412 | XO2-10 | AMaguire | SPC | [5 empty] | 410.5 | 411.1 | 413.5 | 412.7 | 2767.6 | 2767 | 2022-05-10 17:15 | 05/11/2022 08:50 | 22.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.4 | 0.34105 | -0.6 | -0.021679 | 2 | 0.36273 | Excellent | 2022 | 5 | in | in
| VCU20220412 | XO2-32 | AMaguire | SPC | [5 empty] | 411.4 | 411.3 | 413.7 | 413.4 | 2768.4 | 2767 | 2022-05-10 18:15 | 05/11/2022 08:00 | 20.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.048614 | -1.4 | -0.050571 | 1.6 | 0.099185 | Excellent | 2022 | 5 | in | in
| TRB820220418 | XO2-22 | LDavis | SPC | [5 empty] | 91.8 | 92 | 92.9 | 92.9 | 2769.3 | 2767 | 2022-05-11 14:30 | 05/12/2022 07:45 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.21786 | -2.3 | -0.083053 | 2.5 | 0.30092 | Excellent | 2022 | 5 | in | in
| TRTT20220418 | XO2-27 | LDavis | SPC | [5 empty] | 94.1 | 94 | 97.2 | 95.6 | 2769.1 | 2767 | 2022-05-11 15:45 | 05/12/2022 08:11 | 23.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.5 | 1.594 | -2.1 | -0.075837 | 3.6 | 1.6699 | Excellent | 2022 | 5 | in | in
| LAB20220418 | XO2-57 | LDavis | SPC | [5 empty] | 94.1 | 94 | 95.2 | 95.1 | 2767.7 | 2767 | 2022-05-11 16:45 | 05/12/2022 08:25 | 22.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.4211e-14 | 1.5102e-14 | -0.7 | -0.025292 | 0.7 | 0.025292 | Excellent | 2022 | 5 | in | in
| FAR20220418 | XO2-58 | LDavis | SPC | [5 empty] | 94.2 | 94.3 | 95.3 | 95.2 | 2765.2 | 2767 | 2022-05-11 17:45 | 05/12/2022 08:40 | 21.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.21231 | 1.8 | 0.065095 | 2 | 0.27741 | Excellent | 2022 | 5 | in | in
| TRN20220419 | XO2-34 | TSalman | SPC | [5 empty] | 540 | 540 | 540 | 541 | 2757.9 | 2767 | 2022-05-12 16:30 | 05/13/2022 14:10 | 28.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18519 | 9.1 | 0.32996 | 10.1 | 0.51515 | Excellent | 2022 | 5 | in | out
| ORQ20220419 | XO2-71 | TSalman | SPC | [5 empty] | 540 | 540 | 549 | 549 | 2789 | 2767 | 2022-05-12 17:15 | 05/13/2022 15:35 | 29.3 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -22 | -0.78881 | 22 | 0.78881 | Excellent | 2022 | 5 | in | out
| OBI20220419 | XO2-66 | TSalman | SPC | [5 empty] | 539 | 540 | 536 | 544 | 2774.1 | 2767 | 2022-05-12 18:15 | 05/13/2022 15:01 | 27.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -7 | -1.2987 | -7.1 | -0.25594 | 14.1 | 1.5546 | Excellent | 2022 | 5 | in | out
| HOL20220419 | XO2-70 | TSalman | SPC | [5 empty] | 540 | 540 | 545 | 535 | 2776.7 | 2767 | 2022-05-12 19:00 | 05/12/2022 12:11 | 0.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 10 | 1.8519 | -9.7 | -0.34934 | 19.7 | 2.2012 | Excellent | 2022 | 5 | in | out
| HLT20220419 | XO2-43 | TSalman | SPC | [5 empty] | 539 | 540 | 543 | 543 | 2777 | 2767 | 2022-05-12 20:30 | 05/13/2022 15:15 | 25.8 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.18553 | -10 | -0.3601 | 11 | 0.54563 | Excellent | 2022 | 5 | in | out
| MHO20220419 | XO2-40 | BJones | SPC | [5 empty] | 511 | 510 | 512 | 513 | 2767.3 | 2767 | 2022-05-18 15:00 | 05/19/2022 09:21 | 25.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.39139 | -0.3 | -0.010841 | 2.3 | 0.40223 | Excellent | 2022 | 5 | in | in
| MRU20220419 | XO2-24 | BJones | SPC | [5 empty] | 696 | 696 | 699 | 696 | 2764.7 | 2767 | 2022-05-18 15:45 | 05/19/2022 09:36 | 24.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.43103 | 2.3 | 0.083192 | 5.3 | 0.51423 | Excellent | 2022 | 5 | in | in
| TWA20220419 | XO2-46 | BJones | SPC | [5 empty] | 697 | 695 | 692 | 698 | 2768.2 | 2767 | 2022-05-18 16:30 | 05/19/2022 10:12 | 24.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -8 | -1.1478 | -1.2 | -0.043349 | 9.2 | 1.1911 | Excellent | 2022 | 5 | in | in
| PDC20220420 | XO2-74 | DColvin | SPC | [5 empty] | 452.8 | 452.9 | 455.9 | 454.7 | 2804 | 2767 | 2022-05-18 17:30 | 05/19/2022 10:00 | 23.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.3 | 0.2871 | -37 | -1.3195 | 38.3 | 1.6066 | Excellent | 2022 | 5 | in | out
| MRX20220419 | XO2-35 | BJones | SPC | [5 empty] | 378.7 | 379 | 378.5 | 379.6 | 2759.4 | 2767 | 2022-05-18 18:00 | 05/19/2022 10:00 | 23 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.21125 | 7.6 | 0.27542 | 8.4 | 0.48667 | Excellent | 2022 | 5 | in | out
| SGA20220420 | XO2-50 | DColvin | SPC | [5 empty] | 454.1 | 454.3 | 449.8 | 454.3 | 2740 | 2767 | 2022-05-18 18:15 | 05/19/2022 10:00 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4.3 | -0.94693 | 27 | 0.9854 | 31.3 | 1.9323 | Excellent | 2022 | 5 | in | out
| BET20220419 | XO2-65 | JFrantzich | SPC | [5 empty] | 333.6 | 334 | 335.5 | 335.5 | 2772 | 2767 | 2022-05-20 18:00 | 05/20/2022 11:05 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.1199 | -5 | -0.18038 | 5.4 | 0.30028 | Excellent | 2022 | 5 | in | out
| BLP20220419 | XO2-53 | JFrantzich | SPC | [5 empty] | 1269 | 1270 | 1252 | 1270 | 2788 | 2767 | 2022-05-20 19:15 | 05/20/2022 12:20 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -17 | -1.3396 | -21 | -0.75323 | 38 | 2.0929 | Excellent | 2022 | 5 | in | out
| MOK20220426 | XO2-21 | TSalman | SPC | [5 empty] | 147.6 | 147.6 | 148.2 | 146.6 | 2780.4 | 2767 | 2022-05-24 15:30 | 05/25/2022 14:00 | 29.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.6 | 1.084 | -13.4 | -0.48195 | 15 | 1.566 | Excellent | 2022 | 5 | in | out
| OSJ20220426 | XO2-20 | TSalman | SPC | [5 empty] | 148.6 | 148.9 | 152.4 | 149.9 | 2786.7 | 2767 | 2022-05-24 16:15 | 05/25/2022 16:53 | 31.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.8 | 1.8843 | -19.7 | -0.70693 | 22.5 | 2.5912 | Excellent | 2022 | 5 | in | out
| FAL20220426 | XO2-03 | TSalman | SPC | [5 empty] | 151.7 | 152.1 | 153.4 | 153 | 2785.8 | 2767 | 2022-05-24 17:30 | 05/24/2022 10:34 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.52736 | -18.8 | -0.67485 | 19.6 | 1.2022 | Excellent | 2022 | 5 | in | out
| FCT20220426 | XO2-25 | TSalman | SPC | [5 empty] | 153.7 | 153.8 | 155.4 | 152.8 | 2776.8 | 2767 | 2022-05-24 18:00 | 05/25/2022 13:05 | 26.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.7 | 1.7567 | -9.8 | -0.35292 | 12.5 | 2.1096 | Excellent | 2022 | 5 | in | out
| TSL20220426 | XO2-41 | TSalman | SPC | [5 empty] | 156.5 | 156.7 | 157.6 | 157.5 | 2785.4 | 2767 | 2022-05-24 18:30 | 05/25/2022 16:00 | 28.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.19169 | -18.4 | -0.66059 | 18.7 | 0.85228 | Excellent | 2022 | 5 | in | out
| SJD20220503 | XO2-04 | BJones | SPC | [5 empty] | 314 | 314.1 | 311.9 | 311.7 | 2761.5 | 2767 | 2022-05-31 15:15 | 06/01/2022 12:15 | 28 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.095541 | 5.5 | 0.19917 | 5.8 | 0.29471 | Excellent | 2022 | 5 | in | out
| OH120220503 | XO2-49 | BJones | SPC | [5 empty] | 314.6 | 314.7 | 314.3 | 314.1 | 2770 | 2767 | 2022-05-31 15:45 | 06/01/2022 11:30 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.095359 | -3 | -0.1083 | 3.3 | 0.20366 | Excellent | 2022 | 5 | in | out
| ORX20220503 | XO2-05 | BJones | SPC | [5 empty] | 314.2 | 315.3 | 314.5 | 313.7 | 2766.4 | 2767 | 2022-05-31 16:15 | 06/01/2022 11:45 | 26.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.9 | 0.60471 | 0.6 | 0.021689 | 2.5 | 0.6264 | Excellent | 2022 | 5 | in | in
| GLE20220503 | XO2-61 | BJones | SPC | [5 empty] | 318.2 | 317.7 | 316.9 | 316.1 | 2763.5 | 2767 | 2022-05-31 16:45 | 06/01/2022 12:00 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.09428 | 3.5 | 0.12665 | 3.8 | 0.22093 | Excellent | 2022 | 5 | in | out
| ORM20220504 | XO2-02 | AMaguire | SPC | [5 empty] | 443.2 | 442.8 | 443.5 | 443.4 | 2749.3 | 2767 | 2022-06-02 15:30 | 06/03/2022 08:00 | 23.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.06769 | 17.7 | 0.6438 | 18 | 0.71149 | Excellent | 2022 | 6 | in | out
| ODM20220504 | XO2-19 | AMaguire | SPC | [5 empty] | 444.5 | 445.1 | 437.1 | 447 | 2770 | 2767 | 2022-06-02 16:15 | 06/03/2022 08:50 | 23.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -9.3 | -2.0922 | -3 | -0.1083 | 12.3 | 2.2005 | Excellent | 2022 | 6 | in | out
| GLC20220504 | XO2-36 | AMaguire | SPC | [5 empty] | 445.9 | 445.8 | 446.2 | 447 | 2765 | 2767 | 2022-06-02 16:45 | 06/03/2022 09:30 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.9 | -0.20184 | 2 | 0.072333 | 2.9 | 0.27417 | Excellent | 2022 | 6 | in | in
| ORI20220510 | XO2-12 | AMaguire | SPC | [5 empty] | 405.1 | 405.2 | 398.9 | 405.7 | 2750 | 2767 | 2022-06-06 16:15 | 06/06/2022 13:00 | 3.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6.7 | -1.6539 | 17 | 0.61818 | 23.7 | 2.2721 | Excellent | 2022 | 6 | in | out
| WCI20220510 | XO2-42 | AMaguire | SPC | [5 empty] | 405.8 | 406.1 | 406.8 | 406.3 | 2748.5 | 2767 | 2022-06-06 16:45 | 06/06/2022 13:50 | 4.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.19714 | 18.5 | 0.67309 | 19.3 | 0.87024 | Excellent | 2022 | 6 | in | out
| VCU20220510 | XO2-48 | AMaguire | SPC | [5 empty] | 406.6 | 406.9 | 406.1 | 406.9 | 2749.3 | 2767 | 2022-06-06 17:15 | 06/06/2022 13:30 | 3.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.12297 | 17.7 | 0.6438 | 18.2 | 0.76677 | Excellent | 2022 | 6 | in | out
| LIS20220502 | ND-8 | EHuynh | SPC | [5 empty] | 310.8 | 311.2 | 317.5 | 312.3 | 2771.4 | 2767 | 2022-06-06 17:45 | 06/06/2022 10:50 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5.6 | 1.8018 | -4.4 | -0.15876 | 10 | 1.9606 | Excellent | 2022 | 6 | in | out
| TRN20220512 | XO2-60 | TSalman | SPC | [5 empty] | 439.5 | 439.7 | 438.5 | 438.6 | 2798.7 | 2767 | 2022-06-07 16:00 | 06/09/2022 08:20 | 47.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.022753 | -31.7 | -1.1327 | 31.8 | 1.1554 | Excellent | 2022 | 6 | in | out
| OBI20220512 | XO2-32 | TSalman | SPC | [5 empty] | 440.4 | 440.3 | 437.2 | 439.1 | 2790.7 | 2767 | 2022-06-07 17:00 | 06/09/2022 09:05 | 47.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.45413 | -23.7 | -0.84925 | 25.7 | 1.3034 | Excellent | 2022 | 6 | in | out
| HOL20220512 | XO2-11 | TSalman | SPC | [5 empty] | 441.9 | 441.8 | 438.3 | 439.9 | 2770.9 | 2767 | 2022-06-07 18:00 | 06/09/2022 10:25 | 47.4 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.7 | -0.3847 | -3.9 | -0.14075 | 5.6 | 0.52545 | Excellent | 2022 | 6 | in | out
| ORQ20220524 | XO2-66 | TSalman | SPC | [5 empty] | 443.1 | 443.3 | 442.9 | 443.2 | 2774.4 | 2767 | 2022-06-07 19:00 | 06/09/2022 10:45 | 46.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.022568 | -7.4 | -0.26672 | 7.5 | 0.28929 | Excellent | 2022 | 6 | in | out
| HLT20220512 | XO2-56 | TSalman | SPC | [5 empty] | 443.7 | 443.9 | 443.1 | 440.5 | 2760.4 | 2767 | 2022-06-07 19:30 | 06/09/2022 09:55 | 45.4 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.8 | 0.63106 | 6.6 | 0.2391 | 9.4 | 0.87015 | Excellent | 2022 | 6 | in | out
| TRB820220511 | XO2-13 | LDavis | SPC | [5 empty] | 84.9 | 84.8 | 86 | 84.8 | 2765.9 | 2767 | 2022-06-09 14:45 | 06/10/2022 09:23 | 25.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.1 | 1.2956 | 1.1 | 0.03977 | 2.2 | 1.3354 | Excellent | 2022 | 6 | in | in
| TRTT20220511 | XO2-15 | LDavis | SPC | [5 empty] | 87.8 | 87.8 | 88 | 88 | 2767 | 2767 | 2022-06-09 15:45 | 06/10/2022 08:35 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | Excellent | 2022 | 6 | in | in
| LAB20220511 | XO2-16 | LDavis | SPC | [5 empty] | 88.7 | 88.8 | 88.6 | 88.7 | 2762.3 | 2767 | 2022-06-09 16:45 | 06/10/2022 09:22 | 23.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.4211e-14 | -1.6021e-14 | 4.7 | 0.17015 | 4.7 | 0.17015 | Excellent | 2022 | 6 | in | out
| FAR20220511 | XO2-37 | LDavis | SPC | [5 empty] | 89.5 | 89.7 | 90 | 90.1 | 2767.1 | 2767 | 2022-06-09 17:30 | 06/10/2022 09:38 | 23.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.11173 | -0.1 | -0.0036139 | 0.2 | 0.11535 | Excellent | 2022 | 6 | in | in
| SXS20220428 | XO2-44 | PScott | SPC | [5 empty] | 137.9 | 138.4 | 139.3 | 139.6 | 2789 | 2767 | 2022-06-10 16:00 | 06/13/2022 15:25 | 78.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.14503 | -22 | -0.78881 | 22.2 | 0.93385 | Excellent | 2022 | 6 | in | out
| SOI20220428 | XO2-55 | PScott | SPC | [5 empty] | 139 | 139.4 | 144.9 | 140.5 | 2792 | 2767 | 2022-06-10 16:30 | 06/13/2022 14:00 | 76.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4.8 | 3.4532 | -25 | -0.89542 | 29.8 | 4.3487 | Good | 2022 | 6 | in | out
| MIR20220428 | XO2-63 | PScott | SPC | [5 empty] | 135.9 | 135.8 | 136.3 | 136.8 | 2782 | 2767 | 2022-06-10 17:00 | 06/13/2022 13:30 | 75.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.6 | -0.4415 | -15 | -0.53918 | 15.6 | 0.98068 | Excellent | 2022 | 6 | in | out
| PDC20220518 | XO2-73 | DColvin | SPC | [5 empty] | 524 | 525 | 516 | 509 | 2773.6 | 2767 | 2022-06-15 19:45 | 06/16/2022 08:30 | 19.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 8 | 1.5267 | -6.6 | -0.23796 | 14.6 | 1.7647 | Excellent | 2022 | 6 | in | out
| SGA20220518 | XO2-72 | DColvin | SPC | [5 empty] | 437.3 | 437.7 | 434.4 | 436.7 | 2784.6 | 2767 | 2022-06-15 20:30 | 06/16/2022 08:30 | 19 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.9 | -0.43448 | -17.6 | -0.63205 | 19.5 | 1.0665 | Excellent | 2022 | 6 | in | out
| MHO20220518 | XO2-09 | BJones | SPC | [5 empty] | 798 | 799 | 791 | 798 | 2761.3 | 2767 | 2022-06-21 15:30 | 06/22/2022 13:15 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6 | -0.75188 | 5.7 | 0.20642 | 11.7 | 0.9583 | Excellent | 2022 | 6 | in | out
| MRU20220518 | XO2-10 | BJones | SPC | [5 empty] | 798 | 798 | 795 | 797 | 2765.4 | 2767 | 2022-06-21 16:15 | 06/22/2022 14:05 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.25063 | 1.6 | 0.057858 | 3.6 | 0.30848 | Excellent | 2022 | 6 | in | in
| MRX20220518 | XO2-64 | BJones | SPC | [5 empty] | 798 | 798 | 796 | 800 | 2768.2 | 2767 | 2022-06-21 17:00 | 06/22/2022 13:50 | 27.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.50125 | -1.2 | -0.043349 | 5.2 | 0.5446 | Excellent | 2022 | 6 | in | in
| TWA20220518 | XO2-06 | BJones | SPC | [5 empty] | 798 | 798 | 802 | 800 | 2768 | 2767 | 2022-06-21 17:45 | 06/22/2022 13:35 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.25063 | -1 | -0.036127 | 3 | 0.28675 | Excellent | 2022 | 6 | in | in
| BET20220520 | XO2-27 | JFrantzich | SPC | [5 empty] | 647 | 652 | 651 | 651 | 2768.1 | 2767 | 2022-06-21 18:00 | 06/21/2022 11:03 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5 | 0.7728 | -1.1 | -0.039738 | 6.1 | 0.81254 | Excellent | 2022 | 6 | in | in
| BLP20220520 | XO2-34 | JFrantzich | SPC | [5 empty] | 2516 | 2516 | 2500 | 2509 | 2772 | 2767 | 2022-06-21 19:00 | 06/21/2022 16:10 | 4.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -9 | -0.35771 | -5 | -0.18038 | 14 | 0.53809 | Excellent | 2022 | 6 | in | out
| TSL20220524 | XO2-69 | TSalman | SPC | [5 empty] | 2412 | 2412 | 2403 | 2415 | 2771.3 | 2767 | 2022-06-28 15:15 | 06/28/2022 08:25 | 0.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -12 | -0.49751 | -4.3 | -0.15516 | 16.3 | 0.65267 | Excellent | 2022 | 6 | in | out
| FCT20220524 | XO2-65 | TSalman | SPC | [5 empty] | 2408 | 2410 | 2297 | 2411 | 2783.7 | 2767 | 2022-06-28 16:00 | 06/28/2022 09:04 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -112 | -4.6512 | -16.7 | -0.59992 | 128.7 | 5.2511 | Good | 2022 | 6 | in | out
| FAL20220524 | XO2-43 | TSalman | SPC | [5 empty] | 2410 | 2411 | 2403 | 2406 | 2789.8 | 2767 | 2022-06-28 16:30 | 06/28/2022 15:30 | 6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.082988 | -22.8 | -0.81726 | 24.8 | 0.90025 | Excellent | 2022 | 6 | in | out
| OSJ20220524 | XO2-70 | TSalman | SPC | [5 empty] | 673 | 673 | 666 | 679 | 2798.8 | 2767 | 2022-06-28 17:15 | 06/29/2022 13:00 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -13 | -1.9316 | -31.8 | -1.1362 | 44.8 | 3.0679 | Good | 2022 | 6 | in | out
| MOK20220524 | XO2-59 | TSalman | SPC | [5 empty] | 674 | 674 | 675 | 670 | 2776.4 | 2767 | 2022-06-28 17:45 | 06/28/2022 10:53 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5 | 0.74184 | -9.4 | -0.33857 | 14.4 | 1.0804 | Excellent | 2022 | 6 | in | out
| SJD20220531 | XO2-30 | BJones | SPC | [5 empty] | 151.5 | 151.7 | 152.5 | 152.3 | 2773 | 2767 | 2022-06-29 15:00 | 06/30/2022 14:25 | 30.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.26403 | -6 | -0.21637 | 6.4 | 0.4804 | Excellent | 2022 | 6 | in | out
| OH120220531 | XO2-35 | BJones | SPC | [5 empty] | 152.5 | 152.4 | 153.2 | 152.8 | 2767.9 | 2767 | 2022-06-29 15:45 | 06/29/2022 15:00 | 6.2 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.19672 | -0.9 | -0.032516 | 1.2 | 0.22924 | Excellent | 2022 | 6 | in | in
| ORX20220531 | XO2-46 | BJones | SPC | [5 empty] | 153.3 | 153.5 | 153.6 | 153.9 | 2763.3 | 2767 | 2022-06-29 17:00 | 06/29/2022 15:16 | 5.3 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.065232 | 3.7 | 0.1339 | 3.8 | 0.19913 | Excellent | 2022 | 6 | in | out
| GLE20220531 | XO2-24 | AMaguire | SPC | [5 empty] | 446.3 | 446.6 | 446.6 | 448 | 2767.4 | 2767 | 2022-07-05 16:45 | 07/06/2022 08:15 | 22.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.1 | -0.24647 | -0.4 | -0.014454 | 1.5 | 0.26092 | Excellent | 2022 | 7 | in | in
| ORM20220602 | XO2-20 | AMaguire | SPC | [5 empty] | 445.5 | 445.7 | 445.3 | 446.4 | 2770 | 2767 | 2022-07-05 18:00 | 07/06/2022 09:00 | 22 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.9 | -0.20202 | -3 | -0.1083 | 3.9 | 0.31032 | Excellent | 2022 | 7 | in | out
| ODM20220602 | XO2-40 | AMaguire | SPC | [5 empty] | 533 | 533 | 530 | 531 | 2755.7 | 2767 | 2022-07-05 18:45 | 07/06/2022 08:25 | 20.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18762 | 11.3 | 0.41006 | 12.3 | 0.59768 | Excellent | 2022 | 7 | in | out
| GLC20220602 | XO2-53 | AMaguire | SPC | [5 empty] | 534 | 534 | 534 | 534 | 2762 | 2767 | 2022-07-05 19:15 | 07/06/2022 09:00 | 20.8 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 5 | 0.18103 | 5 | 0.18103 | Excellent | 2022 | 7 | in | out
| LIS20220606 | ND-6 | EHuynh | SPC | [5 empty] | 255.7 | 256.3 | 260.6 | 257 | 2775.6 | 2767 | 2022-07-05 20:00 | 07/06/2022 13:06 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4.2 | 1.6425 | -8.6 | -0.30984 | 12.8 | 1.9524 | Excellent | 2022 | 7 | in | out
| LAB20220609 | XO2-57 | LDavis | SPC | [5 empty] | 69.8 | 69.9 | 69.8 | 70 | 2768.7 | 2767 | 2022-07-07 16:00 | 07/07/2022 13:38 | 4.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.14327 | -1.7 | -0.061401 | 1.8 | 0.20467 | Excellent | 2022 | 7 | in | in
| ORI20220606 | XO2-02 | AMaguire | SPC | [5 empty] | 587 | 580 | 580 | 586 | 2765.2 | 2767 | 2022-07-12 16:15 | 07/12/2022 13:45 | 4.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -13 | -2.2147 | 1.8 | 0.065095 | 14.8 | 2.2797 | Excellent | 2022 | 7 | in | in
| WCI20220606 | XO2-19 | AMaguire | SPC | [5 empty] | 588 | 589 | 582 | 588 | 2778.1 | 2767 | 2022-07-12 16:45 | 07/12/2022 13:13 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5 | -0.85034 | -11.1 | -0.39955 | 16.1 | 1.2499 | Excellent | 2022 | 7 | in | out
| VCU20220606 | XO2-36 | AMaguire | SPC | [5 empty] | 590 | 590 | 588 | 589 | 2765.4 | 2767 | 2022-07-12 17:30 | 07/12/2022 14:10 | 3.7 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16949 | 1.6 | 0.057858 | 2.6 | 0.22735 | Excellent | 2022 | 7 | in | in
| WDC20220616 | XO2-14 | CWeber | SPC | [5 empty] | 597 | 598 | 592 | 597 | 2755.3 | 2767 | 2022-07-13 15:15 | 07/14/2022 11:00 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.67002 | 11.7 | 0.42464 | 15.7 | 1.0947 | Excellent | 2022 | 7 | in | out
| TRN20220607 | XO2-49 | TSalman | SPC | [5 empty] | 469.8 | 469.9 | 465.8 | 471.2 | 2779.4 | 2767 | 2022-07-13 15:30 | 07/14/2022 14:20 | 29.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5.3 | -1.1281 | -12.4 | -0.44614 | 17.7 | 1.5743 | Excellent | 2022 | 7 | in | out
| RSCC20220616 | XO2-74 | CWeber | SPC | [5 empty] | 600 | 601 | 595 | 596 | 2744.8 | 2767 | 2022-07-13 16:00 | 07/14/2022 09:20 | 24.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 22.2 | 0.8088 | 22.2 | 0.8088 | Excellent | 2022 | 7 | in | out
| DRB20220616 | XO2-50 | CWeber | SPC | [5 empty] | 598 | 599 | 605 | 605 | 2754.5 | 2767 | 2022-07-13 16:15 | 07/15/2022 13:00 | 51.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.16722 | 12.5 | 0.4538 | 13.5 | 0.62103 | Excellent | 2022 | 7 | in | out
| HLT20220607 | XO2-03 | TSalman | SPC | [5 empty] | 470.3 | 470.3 | 437.2 | 470.6 | 2776.8 | 2767 | 2022-07-13 16:30 | 07/14/2022 13:05 | 27.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -33.4 | -7.1018 | -9.8 | -0.35292 | 43.2 | 7.4548 | Good | 2022 | 7 | in | out
| HCHM20220616 | XO2-04 | CWeber | SPC | [5 empty] | 600 | 600 | 591 | 596 | 2754.7 | 2767 | 2022-07-13 17:00 | 07/14/2022 08:45 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5 | -0.83333 | 12.3 | 0.44651 | 17.3 | 1.2798 | Excellent | 2022 | 7 | in | out
| ORQ20220607 | XO2-21 | TSalman | SPC | [5 empty] | 470.3 | 470.9 | 471.5 | 471.5 | 2773.5 | 2767 | 2022-07-13 17:15 | 07/14/2022 16:10 | 29.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.12758 | -6.5 | -0.23436 | 7.1 | 0.36194 | Excellent | 2022 | 7 | in | out
| HOL20220607 | XO2-41 | TSalman | SPC | [5 empty] | 472 | 472.1 | 473 | 471.8 | 2770.6 | 2767 | 2022-07-13 17:45 | 07/14/2022 13:55 | 27.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.3 | 0.27542 | -3.6 | -0.12994 | 4.9 | 0.40536 | Excellent | 2022 | 7 | in | out
| OBI20220607 | XO2-61 | TSalman | SPC | [5 empty] | 474.7 | 474.7 | 475.9 | 474.7 | 2783.1 | 2767 | 2022-07-13 18:45 | 07/14/2022 13:25 | 25.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.2 | 0.25279 | -16.1 | -0.57849 | 17.3 | 0.83128 | Excellent | 2022 | 7 | in | out
| TRB820220609 | XO2-22 | LDavis | SPC | [5 empty] | 95.8 | 96 | 96.4 | 96.1 | 2771.3 | 2767 | 2022-07-20 14:45 | 07/25/2022 09:03 | 121.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.52192 | -4.3 | -0.15516 | 4.8 | 0.67708 | Excellent | 2022 | 7 | in | out
| TRTT20220609 | XO2-26 | LDavis | SPC | [5 empty] | 99.9 | 100 | 100.1 | 100.1 | 2771.2 | 2767 | 2022-07-20 16:00 | 07/25/2022 09:15 | 120.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.1001 | -4.2 | -0.15156 | 4.3 | 0.25166 | Excellent | 2022 | 7 | in | out
| FAR20220609 | XO2-58 | LDavis | SPC | [5 empty] | 100.9 | 100.9 | 101.3 | 101 | 2771.4 | 2767 | 2022-07-20 17:00 | 07/25/2022 10:04 | 120.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.29732 | -4.4 | -0.15876 | 4.7 | 0.45609 | Excellent | 2022 | 7 | in | out
| SGA20220615 | XO2-08 | DColvin | SPC | [5 empty] | 958 | 959 | 938 | 955 | 2777 | 2767 | 2022-07-21 19:15 | 07/22/2022 15:00 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -16 | -1.6701 | -10 | -0.3601 | 26 | 2.0302 | Excellent | 2022 | 7 | in | out
| PDC20220615 | XO2-47 | DColvin | SPC | [5 empty] | 967 | 966 | 965 | 964 | 2776 | 2767 | 2022-07-21 21:00 | 07/22/2022 12:00 | 22 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -9 | -0.32421 | 9 | 0.32421 | Excellent | 2022 | 7 | in | out
| TSL20220628 | XO2-11 | TSalman | SPC | [5 empty] | 1159 | 1159 | 1157 | 1154 | 2751.5 | 2767 | 2022-07-26 15:00 | 07/27/2022 10:15 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.25884 | 15.5 | 0.56333 | 18.5 | 0.82217 | Excellent | 2022 | 7 | in | out
| FCT20220628 | XO2-27 | TSalman | SPC | [5 empty] | 1156 | 1156 | 1143 | 1154 | 2789 | 2767 | 2022-07-26 15:45 | 07/26/2022 14:00 | 5.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -11 | -0.95156 | -22 | -0.78881 | 33 | 1.7404 | Excellent | 2022 | 7 | in | out
| FAL20220628 | XO2-56 | TSalman | SPC | [5 empty] | 1156 | 1156 | 1151 | 1142 | 2782 | 2767 | 2022-07-26 16:15 | 07/27/2022 09:16 | 24 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 9 | 0.77855 | -15 | -0.53918 | 24 | 1.3177 | Excellent | 2022 | 7 | in | out
| OSJ20220628 | XO2-71 | TSalman | SPC | [5 empty] | 1153 | 1154 | 1145 | 1133 | 2754.8 | 2767 | 2022-07-26 16:45 | 07/27/2022 10:45 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 13 | 1.1275 | 12.2 | 0.44286 | 25.2 | 1.5704 | Excellent | 2022 | 7 | in | out
| MOK20220628 | XO2-18 | TSalman | SPC | [5 empty] | 1150 | 1153 | 1122 | 1152 | 2787 | 2767 | 2022-07-26 17:15 | 07/26/2022 13:45 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -27 | -2.3478 | -20 | -0.71762 | 47 | 3.0654 | Good | 2022 | 7 | in | out
| BLP20220621 | XO2-32 | JFrantzich | SPC | [5 empty] | 955 | 956 | 951 | 952 | 2771.5 | 2767 | 2022-07-26 18:30 | 07/27/2022 15:31 | 28 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -4.5 | -0.16237 | 4.5 | 0.16237 | Excellent | 2022 | 7 | in | out
| BET20220621 | XO2-66 | JFrantzich | SPC | [5 empty] | 817 | 817 | 816 | 816 | 2779 | 2767 | 2022-07-26 19:30 | 07/26/2022 15:00 | 2.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -12 | -0.43181 | 12 | 0.43181 | Excellent | 2022 | 7 | in | out
| MHO20220621 | XO2-60 | BJones | SPC | [5 empty] | 689 | 689 | 679 | 663 | 2763.5 | 2767 | 2022-07-28 15:45 | 07/29/2022 11:35 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 16 | 2.3222 | 3.5 | 0.12665 | 19.5 | 2.4489 | Excellent | 2022 | 7 | in | out
| SXS20220610 | XO2-12 | PScott | SPC | [5 empty] | 114.6 | 114.7 | 114.6 | 114.6 | 2756.6 | 2767 | 2022-07-28 16:30 | 07/29/2022 12:00 | 26.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.08726 | 10.4 | 0.37728 | 10.5 | 0.46454 | Excellent | 2022 | 7 | in | out
| MRU20220621 | XO2-44 | BJones | SPC | [5 empty] | 688 | 688 | 689 | 384.6 | 2760.9 | 2767 | 2022-07-28 16:30 | 07/29/2022 10:50 | 25.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 304.4 | 44.2442 | 6.1 | 0.22094 | 310.5 | 44.4651 | MAL | 2022 | 7 | in | out
| SOI20220610 | XO2-42 | PScott | SPC | [5 empty] | 115.3 | 115.4 | 116.8 | 114.6 | 2743.1 | 2767 | 2022-07-28 17:00 | 07/29/2022 12:00 | 26 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.3 | 1.9948 | 23.9 | 0.87128 | 26.2 | 2.8661 | Excellent | 2022 | 7 | in | out
| MRX20220621 | XO2-37 | BJones | SPC | [5 empty] | 685 | 687 | 683 | 674 | 2764.7 | 2767 | 2022-07-28 17:30 | 07/29/2022 11:05 | 24.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 11 | 1.6058 | 2.3 | 0.083192 | 13.3 | 1.689 | Excellent | 2022 | 7 | in | in
| MIR20220610 | XO2-48 | PScott | SPC | [5 empty] | 109.8 | 110.1 | 110.1 | 110.2 | 2771.5 | 2767 | 2022-07-28 17:30 | 07/29/2022 12:00 | 25.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.18215 | -4.5 | -0.16237 | 4.7 | 0.34452 | Excellent | 2022 | 7 | in | out
| TWA20220621 | XO2-63 | BJones | SPC | [5 empty] | 686 | 684 | 684 | 676 | 2763.9 | 2767 | 2022-07-28 18:15 | 07/29/2022 11:20 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 6 | 0.87464 | 3.1 | 0.11216 | 9.1 | 0.9868 | Excellent | 2022 | 7 | in | out
| SJD20220629 | XO2-73 | BJones | SPC | [5 empty] | 453.1 | 453.5 | 447.1 | 452.7 | 2761.6 | 2767 | 2022-08-02 15:15 | 08/02/2022 14:25 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5.2 | -1.1476 | 5.4 | 0.19554 | 10.6 | 1.3432 | Excellent | 2022 | 8 | in | out
| OH120220629 | XO2-10 | BJones | SPC | [5 empty] | 453.9 | 454.3 | 453 | 453 | 2761.3 | 2767 | 2022-08-02 15:30 | 08/02/2022 14:06 | 5.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.088125 | 5.7 | 0.20642 | 6.1 | 0.29455 | Excellent | 2022 | 8 | in | out
| ORX20220629 | XO2-34 | BJones | SPC | [5 empty] | 456.7 | 456.6 | 455.6 | 456 | 2763 | 2767 | 2022-08-02 16:30 | 08/02/2022 13:53 | 4.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.10948 | 4 | 0.14477 | 4.5 | 0.25425 | Excellent | 2022 | 8 | in | out
| GLE20220705 | XO2-64 | BJones | SPC | [5 empty] | 459 | 459.6 | 454.9 | 458.3 | 2760 | 2767 | 2022-08-02 17:00 | 08/02/2022 13:37 | 3.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2.8 | -0.61002 | 7 | 0.25362 | 9.8 | 0.86364 | Excellent | 2022 | 8 | in | out
| LIS20220705 | ND-4 | EHuynh | SPC | [5 empty] | 211.3 | 211.6 | 212.7 | 213.8 | 2768.4 | 2767 | 2022-08-02 17:15 | 08/02/2022 13:10 | 2.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.8 | -0.37861 | -1.4 | -0.050571 | 2.2 | 0.42918 | Excellent | 2022 | 8 | in | in
| GLC20220705 | XO2-06 | AMaguire | SPC | [5 empty] | 165.8 | 165.4 | 165.9 | 163.9 | 2765.2 | 2767 | 2022-08-03 16:03 | 08/04/2022 16:00 | 30.9 | NCRO | FALSE | FALSE | TRUE |  | TRUE |  |  |  | 1.6 | 0.96502 | 1.8 | 0.065095 | 3.4 | 1.0301 | Excellent | 2022 | 8 | in | in
| ODM20220705 | XO2-35 | AMaguire | SPC | [5 empty] | 460.6 | 461.1 | 459.3 | 459.3 | 2762.8 | 2767 | 2022-08-03 16:30 | 08/03/2022 15:00 | 5.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.10855 | 4.2 | 0.15202 | 4.7 | 0.26057 | Excellent | 2022 | 8 | in | out
| ORM20220705 | XO2-46 | AMaguire | SPC | [5 empty] | 462.4 | 462.3 | 461.3 | 460.6 | 2767.3 | 2767 | 2022-08-03 17:00 | 08/03/2022 15:20 | 5.3 | NCRO |  | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.12976 | -0.3 | -0.010841 | 0.9 | 0.1406 | Excellent | 2022 | 8 | in | in
| STTD20220727 | ND-1 | EHuynh | SPC | [5 empty] | 164.5 | 165 | 165 | 165.1 | 2765.5 | 2767 | 2022-08-08 17:45 | 08/08/2022 14:15 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.4 | 0.24316 | 1.5 | 0.05424 | 1.9 | 0.2974 | Excellent | 2022 | 8 | in | in
| LIS20220802 | ND-3 | EHuynh | SPC | [5 empty] | 200 | 200.6 | 199.7 | 198.9 | 2761.7 | 2767 | 2022-08-08 18:45 | 08/08/2022 15:17 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.4 | 0.7 | 5.3 | 0.19191 | 6.7 | 0.89191 | Excellent | 2022 | 8 | in | out
| RD2220220727 | ND-8 | EHuynh | SPC | [5 empty] | 906 | 906 | 907 | 907 | 2769.5 | 2767 | 2022-08-09 16:00 | 08/09/2022 16:07 | 7.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -2.5 | -0.090269 | 2.5 | 0.090269 | Excellent | 2022 | 8 | in | in
| VCU20220712 | XO2-53 | AMaguire | SPC | [5 empty] | 391.7 | 391.6 | 391.1 | 391.1 | 2757 | 2767 | 2022-08-09 16:30 | 08/10/2022 09:30 | 24 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.02553 | 10 | 0.36271 | 10.1 | 0.38824 | Excellent | 2022 | 8 | in | out
| WCI20220712 | XO2-24 | AMaguire | SPC | [5 empty] | 392 | 392.4 | 369.1 | 391.8 | 2759 | 2767 | 2022-08-09 17:00 | 08/10/2022 09:45 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -22.3 | -5.6888 | 8 | 0.28996 | 30.3 | 5.9787 | Good | 2022 | 8 | in | out
| ORI20220712 | XO2-20 | AMaguire | SPC | [5 empty] | 393.5 | 394 | 378.89 | 391.6 | 2760 | 2767 | 2022-08-09 17:30 | 08/10/2022 10:08 | 23.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -12.21 | -3.1029 | 7 | 0.25362 | 19.21 | 3.3565 | Good | 2022 | 8 | in | out
| I8020220727 | ND-5 | EHuynh | SPC | [5 empty] | 506 | 507 | 506 | 506 | 2762.3 | 2767 | 2022-08-09 17:45 | 08/09/2022 11:28 | 0.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.19763 | 4.7 | 0.17015 | 5.7 | 0.36778 | Excellent | 2022 | 8 | in | out
| LAB20220707 | XO2-13 | LDavis | SPC | [5 empty] | 72 | 72.2 | 71.6 | 71.7 | 2769.9 | 2767 | 2022-08-10 16:45 | 08/11/2022 10:39 | 24.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.13889 | -2.9 | -0.1047 | 3 | 0.24359 | Excellent | 2022 | 8 | in | in
| FAR20220720 | XO2-65 | LDavis | SPC | [5 empty] | 74.1 | 74.1 | 74 | 73.9 | 2770.6 | 2767 | 2022-08-10 17:45 | 08/11/2022 10:50 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.13495 | -3.6 | -0.12994 | 3.7 | 0.26489 | Excellent | 2022 | 8 | in | out
| TRTT20220720 | XO2-61 | LDavis | SPC | [5 empty] | 75 | 75 | 74.3 | 74.1 | 2770.9 | 2767 | 2022-08-10 19:15 | 08/11/2022 10:22 | 22.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.26667 | -3.9 | -0.14075 | 4.1 | 0.40742 | Excellent | 2022 | 8 | in | out
| TRB820220720 | XO2-41 | LDavis | SPC | [5 empty] | 75.6 | 75.8 | 75.4 | 75.5 | 2769.1 | 2767 | 2022-08-10 20:15 | 08/11/2022 09:07 | 19.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.13228 | -2.1 | -0.075837 | 2.2 | 0.20811 | Excellent | 2022 | 8 | in | in
| MRX20220728 | XO2-36 | BJones | SPC | [5 empty] | 733 | 733 | 729 | 729 | 2774 | 2767 | 2022-08-16 14:45 | 08/17/2022 13:00 | 29.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -7 | -0.25234 | 7 | 0.25234 | Excellent | 2022 | 8 | in | out
| MRU20220728 | XO2-02 | BJones | SPC | [5 empty] | 733 | 733 | 733 | 734 | 2775.2 | 2767 | 2022-08-16 15:30 | 08/17/2022 13:20 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.13643 | -8.2 | -0.29547 | 9.2 | 0.4319 | Excellent | 2022 | 8 | in | out
| WDC20220713 | XO2-30 | CWeber | SPC | [5 empty] | 438.9 | 439.5 | 436.1 | 441.8 | 2772.2 | 2767 | 2022-08-16 16:15 | 08/17/2022 10:00 | 24.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5.1 | -1.162 | -5.2 | -0.18758 | 10.3 | 1.3496 | Excellent | 2022 | 8 | in | out
| RSCC20220713 | XO2-72 | CWeber | SPC | [5 empty] | 440.6 | 439.5 | 442 | 442.3 | 2773.6 | 2767 | 2022-08-16 16:45 | 08/17/2022 10:00 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.4 | -0.31775 | -6.6 | -0.23796 | 8 | 0.55571 | Excellent | 2022 | 8 | in | out
| MHO20220728 | XO2-16 | BJones | SPC | [5 empty] | 733 | 733 | 731 | 732 | 2774.3 | 2767 | 2022-08-16 16:45 | 08/17/2022 13:40 | 27.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.13643 | -7.3 | -0.26313 | 8.3 | 0.39956 | Excellent | 2022 | 8 | in | out
| DRB20220713 | XO2-45 | CWeber | SPC | [5 empty] | 440.4 | 439.4 | 440.8 | 441.3 | 2776.3 | 2767 | 2022-08-16 17:15 | 08/17/2022 10:00 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.5 | -0.3406 | -9.3 | -0.33498 | 10.8 | 0.67558 | Excellent | 2022 | 8 | in | out
| TWA20220728 | XO2-56 | BJones | SPC | [5 empty] | 733 | 733 | 737 | 736 | 2779.7 | 2767 | 2022-08-16 17:30 | 08/17/2022 12:40 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.13643 | -12.7 | -0.45688 | 13.7 | 0.59331 | Excellent | 2022 | 8 | in | out
| HCHM20220713 | XO2-23 | CWeber | SPC | [5 empty] | 440.2 | 439.2 | 437.4 | 437.4 | 2773.6 | 2767 | 2022-08-16 17:45 | 08/17/2022 10:00 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.22717 | -6.6 | -0.23796 | 7.6 | 0.46513 | Excellent | 2022 | 8 | in | out
| SGA20220721 | XO2-14 | DColvin | SPC | [5 empty] | 1034 | 1034 | 1036 | 1038 | 2787 | 2767 | 2022-08-17 16:45 | 08/18/2022 13:27 | 27.7 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.19342 | -20 | -0.71762 | 22 | 0.91104 | Excellent | 2022 | 8 | in | out
| PDC20220721 | XO2-74 | DColvin | SPC | [5 empty] | 1038 | 1037 | 1034 | 1031 | 2771.6 | 2767 | 2022-08-17 17:00 | 08/18/2022 13:32 | 27.5 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.19268 | -4.6 | -0.16597 | 6.6 | 0.35865 | Excellent | 2022 | 8 | in | out
| SXS20220728 | XO2-15 | PScott | SPC | [5 empty] | 123.1 | 123.3 | 124.5 | 124.4 | 2768.4 | 2767 | 2022-08-18 16:15 | 08/19/2022 11:00 | 25.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.2437 | -1.4 | -0.050571 | 1.7 | 0.29428 | Excellent | 2022 | 8 | in | in
| SOI20220728 | XO2-49 | PScott | SPC | [5 empty] | 124.1 | 124.1 | 125.4 | 125.4 | 2768.7 | 2767 | 2022-08-18 17:00 | 08/19/2022 11:00 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -1.7 | -0.061401 | 1.7 | 0.061401 | Excellent | 2022 | 8 | in | in
| MIR20220728 | XO2-57 | PScott | SPC | [5 empty] | 124.8 | 125 | 125.4 | 125.5 | 2761.2 | 2767 | 2022-08-18 17:45 | 08/19/2022 11:00 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.080128 | 5.8 | 0.21005 | 5.9 | 0.29018 | Excellent | 2022 | 8 | in | out
| TRN20220713 | XO2-55 | TSalman | SPC | [5 empty] | 357.9 | 357.9 | 351.5 | 358.2 | 2780.7 | 2767 | 2022-08-23 15:30 | 08/24/2022 13:50 | 29.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6.7 | -1.872 | -13.7 | -0.49268 | 20.4 | 2.3647 | Excellent | 2022 | 8 | in | out
| HLT20220713 | XO2-43 | TSalman | SPC | [5 empty] | 358.3 | 358.2 | 353.6 | 355.2 | 2767.2 | 2767 | 2022-08-23 16:15 | 08/24/2022 16:10 | 30.9 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.7 | -0.47446 | -0.2 | -0.0072275 | 1.9 | 0.48169 | Excellent | 2022 | 8 | in | in
| ORQ20220713 | XO2-54 | TSalman | SPC | [5 empty] | 359 | 360.6 | 358 | 359.1 | 2773.2 | 2767 | 2022-08-23 17:00 | 08/24/2022 15:53 | 29.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.13928 | -6.2 | -0.22357 | 6.7 | 0.36284 | Excellent | 2022 | 8 | in | out
| HOL20220713 | XO2-69 | TSalman | SPC | [5 empty] | 360.4 | 360.7 | 360.7 | 357.7 | 2771.4 | 2767 | 2022-08-23 17:30 | 08/24/2022 15:30 | 29 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3.3 | 0.91565 | -4.4 | -0.15876 | 7.7 | 1.0744 | Excellent | 2022 | 8 | in | out
| OBI20220713 | XO2-09 | TSalman | SPC | [5 empty] | 362.1 | 362.4 | 358.1 | 361.7 | 2775.5 | 2767 | 2022-08-23 18:15 | 08/24/2022 13:25 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3.3 | -0.91135 | -8.5 | -0.30625 | 11.8 | 1.2176 | Excellent | 2022 | 8 | in | out
| TSL20220726 | XO2-21 | TSalman | SPC | [5 empty] | 1468 | 1467 | 1438 | 1466 | 2771.5 | 2767 | 2022-08-31 15:15 | 09/01/2022 12:40 | 28.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -29 | -1.9755 | -4.5 | -0.16237 | 33.5 | 2.1378 | Excellent | 2022 | 8 | in | out
| FCT20220726 | XO2-25 | TSalman | SPC | [5 empty] | 1465 | 1465 | 1461 | 1466 | 2773.7 | 2767 | 2022-08-31 16:00 | 09/01/2022 13:10 | 28.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5 | -0.3413 | -6.7 | -0.24155 | 11.7 | 0.58285 | Excellent | 2022 | 8 | in | out
| FAL20220726 | XO2-03 | TSalman | SPC | [5 empty] | 1464 | 1465 | 1458 | 1455 | 2770.4 | 2767 | 2022-08-31 16:30 | 09/01/2022 11:45 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4 | 0.27322 | -3.4 | -0.12273 | 7.4 | 0.39595 | Excellent | 2022 | 8 | in | out
| OSJ20220726 | XO2-75 | TSalman | SPC | [5 empty] | 1462 | 1462 | 1432 | 1462 | 2765.2 | 2767 | 2022-08-31 17:15 | 08/31/2022 10:24 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -30 | -2.052 | 1.8 | 0.065095 | 31.8 | 2.1171 | Excellent | 2022 | 8 | in | in
| MOK20220726 | XO2-28 | TSalman | SPC | [5 empty] | 1457 | 1462 | 1441 | 1456 | 2769 | 2767 | 2022-08-31 18:00 | 09/01/2022 13:30 | 26.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -10 | -0.68634 | -2 | -0.072228 | 12 | 0.75857 | Excellent | 2022 | 8 | in | in
| BET20220726 | XO2-26 | JFrantzich | SPC | [5 empty] | 799 | 798 | 802 | 801 | 2777 | 2767 | 2022-09-01 17:45 | 09/02/2022 10:54 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -10 | -0.3601 | 10 | 0.3601 | Excellent | 2022 | 9 | in | out
| BLP20220726 | XO2-19 | JFrantzich | SPC | [5 empty] | 1776 | 1777 | 1778 | 1778 | 2766 | 2767 | 2022-09-01 19:00 | 09/01/2022 12:02 | 0 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.056306 | 1 | 0.036153 | 2 | 0.09246 | Excellent | 2022 | 9 | in | in
| STTD20220808 | ND-6 | EHuynh | SPC | [5 empty] | 181.2 | 181.5 | 181.7 | 180.8 | 2771.5 | 2767 | 2022-09-06 14:30 | 09/06/2022 11:38 | 4.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.2 | 0.66225 | -4.5 | -0.16237 | 5.7 | 0.82462 | Excellent | 2022 | 9 | in | out
| LIS20220808 | ND-4 | EHuynh | SPC | [5 empty] | 265.7 | 266 | 267.4 | 266.4 | 2780.1 | 2767 | 2022-09-06 15:30 | 09/06/2022 08:35 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.3 | 0.48927 | -13.1 | -0.47121 | 14.4 | 0.96048 | Excellent | 2022 | 9 | in | out
| RD2220220809 | ND-1 | EHuynh | SPC | [5 empty] | 863 | 862 | 862 | 862 | 2779.4 | 2767 | 2022-09-07 13:00 | 09/08/2022 11:25 | 29.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.11587 | -12.4 | -0.44614 | 13.4 | 0.56201 | Excellent | 2022 | 9 | in | out
| SJD20220802 | XO2-31 | BJones | SPC | [5 empty] | 603 | 603 | 603 | 603 | 2772.4 | 2767 | 2022-09-07 13:45 | 09/08/2022 11:05 | 28.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -5.4 | -0.19478 | 5.4 | 0.19478 | Excellent | 2022 | 9 | in | out
| OH120220802 | XO2-44 | BJones | SPC | [5 empty] | 603 | 603 | 605 | 605 | 2770 | 2767 | 2022-09-07 14:00 | 09/08/2022 10:40 | 27.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -3 | -0.1083 | 3 | 0.1083 | Excellent | 2022 | 9 | in | out
| I8020220809 | ND-3 | EHuynh | SPC | [5 empty] | 645 | 645 | 644 | 644 | 2779.3 | 2767 | 2022-09-07 15:00 | 09/07/2022 11:52 | 3.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -12.3 | -0.44256 | 12.3 | 0.44256 | Excellent | 2022 | 9 | in | out
| ORX20220802 | XO2-63 | BJones | SPC | [5 empty] | 604 | 605 | 605 | 607 | 2763.8 | 2767 | 2022-09-07 15:00 | 09/08/2022 10:55 | 26.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16556 | 3.2 | 0.11578 | 4.2 | 0.28135 | Excellent | 2022 | 9 | in | out
| OSJ20220831 | XO2-59 | TSalman | SPC | [5 empty] | 168.5 | 168.7 | 169.9 | 169.3 | 2774.6 | 2767 | 2022-09-08 15:00 | 09/08/2022 15:00 | 7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.47478 | -7.6 | -0.27391 | 8.4 | 0.74869 | Excellent | 2022 | 9 | in | out
| HLT20220823 | XO2-71 | TSalman | SPC | [5 empty] | 168.4 | 168.5 | 169.3 | 169.3 | 2776.2 | 2767 | 2022-09-08 16:45 | 09/08/2022 14:25 | 4.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.059382 | -9.2 | -0.33139 | 9.3 | 0.39077 | Excellent | 2022 | 9 | in | out
| TRN20220823 | XO2-49 | TSalman | SPC | [5 empty] | 169.3 | 168.1 | 169.3 | 167.8 | 2775.9 | 2767 | 2022-09-08 17:30 | 09/08/2022 10:32 | 0 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.3 | 0.1772 | -8.9 | -0.32062 | 9.2 | 0.49782 | Excellent | 2022 | 9 | in | out
| GLE20220802 | XO2-60 | BJones | SPC | [5 empty] | 848 | 848 | 847 | 847 | 2766 | 2767 | 2022-09-13 16:15 | 09/14/2022 16:15 | 31 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 1 | 0.036153 | 1 | 0.036153 | Excellent | 2022 | 9 | in | in
| GLC20220803 | XO2-12 | AMaguire | SPC | [5 empty] | 847 | 847 | 844 | 844 | 2772.5 | 2767 | 2022-09-13 17:15 | 09/14/2022 12:00 | 25.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | -5.5 | -0.19838 | 5.5 | 0.19838 | Excellent | 2022 | 9 | in | out
| ORM20220803 | XO2-42 | AMaguire | SPC | [5 empty] | 474.7 | 474.9 | 474.5 | 474.9 | 2766.3 | 2767 | 2022-09-13 18:00 | 09/14/2022 12:25 | 25.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.042132 | 0.7 | 0.025305 | 0.9 | 0.067436 | Excellent | 2022 | 9 | in | in
| ODM20220803 | XO2-37 | AMaguire | SPC | [5 empty] | 475.4 | 475.5 | 475.8 | 475.8 | 2769.5 | 2767 | 2022-09-13 18:45 | 09/14/2022 13:10 | 25.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.021035 | -2.5 | -0.090269 | 2.6 | 0.1113 | Excellent | 2022 | 9 | in | in
| ORI20220809 | XO2-48 | AMaguire | SPC | [5 empty] | 475.9 | 475.7 | 474.9 | 476.3 | 2771.6 | 2767 | 2022-09-13 19:15 | 09/14/2022 11:30 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.6 | -0.33621 | -4.6 | -0.16597 | 6.2 | 0.50217 | Excellent | 2022 | 9 | in | out
| WCI20220809 | XO2-34 | AMaguire | SPC | [5 empty] | 476.3 | 476.4 | 464.2 | 476.7 | 2769.3 | 2767 | 2022-09-13 19:45 | 09/14/2022 11:00 | 22.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -12.4 | -2.6034 | -2.3 | -0.083053 | 14.7 | 2.6865 | Excellent | 2022 | 9 | in | in
| VCU20220809 | XO2-10 | AMaguire | SPC | [5 empty] | 394.5 | 477 | 395 | 477.9 | 2774.9 | 2767 | 2022-09-13 20:15 | 09/14/2022 12:45 | 23.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.4 | -0.10139 | -7.9 | -0.28469 | 8.3 | 0.38609 | Excellent | 2022 | 9 | in | out
| TRB820220810 | XO2-11 | LDavis | SPC | [5 empty] | 104.1 | 104.3 | 104.5 | 104.8 | 2763 | 2767 | 2022-09-14 14:45 | 09/15/2022 10:14 | 26.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.096061 | 4 | 0.14477 | 4.1 | 0.24083 | Excellent | 2022 | 9 | in | out
| MHO20220816 | XO2-32 | BJones | SPC | [5 empty] | 637 | 638 | 638 | 636 | 2767.7 | 2767 | 2022-09-14 15:15 | 09/15/2022 15:00 | 30.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.47096 | -0.7 | -0.025292 | 3.7 | 0.49625 | Excellent | 2022 | 9 | in | in
| TRTT20220810 | XO2-22 | LDavis | SPC | [5 empty] | 106.4 | 106.5 | 106.7 | 106.7 | 2757.2 | 2767 | 2022-09-14 15:45 | 09/15/2022 10:29 | 25.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.093985 | 9.8 | 0.35543 | 9.9 | 0.44942 | Excellent | 2022 | 9 | in | out
| MRU20220816 | XO2-76 | BJones | SPC | [5 empty] | 637 | 638 | 637 | 637 | 2759.9 | 2767 | 2022-09-14 16:00 | 09/14/2022 15:40 | 6.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.15699 | 7.1 | 0.25726 | 8.1 | 0.41424 | Excellent | 2022 | 9 | in | out
| LAB20220810 | XO2-27 | LDavis | SPC | [5 empty] | 107 | 107.1 | 107.2 | 107.2 | 2766.7 | 2767 | 2022-09-14 16:45 | 09/15/2022 10:45 | 25 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.093458 | 0.3 | 0.010843 | 0.4 | 0.1043 | Excellent | 2022 | 9 | in | in
| MRX20220816 | XO2-46 | BJones | SPC | [5 empty] | 638 | 637 | 635 | 636 | 2755.1 | 2767 | 2022-09-14 16:45 | 09/14/2022 16:00 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.31348 | 11.9 | 0.43193 | 13.9 | 0.74541 | Excellent | 2022 | 9 | in | out
| TWA20220816 | XO2-24 | BJones | SPC | [5 empty] | 638 | 637 | 637 | 636 | 2758.5 | 2767 | 2022-09-14 17:15 | 09/15/2022 15:35 | 29.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 8.5 | 0.30814 | 8.5 | 0.30814 | Excellent | 2022 | 9 | in | out
| FAR20220810 | XO2-58 |  | SPC | [5 empty] | 107.5 | 107.8 | 108.8 | 107.9 | 2765.9 | 2767 | 2022-09-14 17:45 | 09/15/2022 11:02 | 24.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.2 | 1.1163 | 1.1 | 0.03977 | 2.3 | 1.156 | Excellent | 2022 | 9 | in | in
| WDC20220816 | XO2-47 | Cweber | SPC | [5 empty] | 486.9 | 487 | 487.8 | 489.8 | 2781.3 | 2767 | 2022-09-20 15:30 | 09/21/2022 15:00 | 30.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.9 | -0.39022 | -14.3 | -0.51415 | 16.2 | 0.90437 | Excellent | 2022 | 9 | in | out
| SXS20220818 | XO2-20 | PScott | SPC | [5 empty] | 130.3 | 130.4 | 130.3 | 130.2 | 2765.1 | 2767 | 2022-09-20 16:00 | 09/21/2022 15:00 | 30 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.15349 | 1.9 | 0.068714 | 2.1 | 0.22221 | Excellent | 2022 | 9 | in | in
| RSCC20220816 | XO2-73 | PScott | SPC | [5 empty] | 487.8 | 488.7 | 489.6 | 488.1 | 2768.1 | 2767 | 2022-09-20 16:15 | 09/21/2022 15:00 | 29.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2.4 | 0.492 | -1.1 | -0.039738 | 3.5 | 0.53174 | Excellent | 2022 | 9 | in | in
| SOI20220818 | XO2-36 |  | SPC | [5 empty] | 130.7 | 130.9 | 131.6 | 131.1 | 2765.7 | 2767 | 2022-09-20 16:30 | 09/21/2022 15:00 | 29.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.7 | 0.53558 | 1.3 | 0.047004 | 2 | 0.58258 | Excellent | 2022 | 9 | in | in
| DRB20220816 | XO2-50 | CWeber | SPC | [5 empty] | 488.8 | 489.5 | 497.1 | 498.3 | 2798.8 | 2767 | 2022-09-20 16:45 | 09/21/2022 15:00 | 29.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.5 | -0.10229 | -31.8 | -1.1362 | 32.3 | 1.2385 | Excellent | 2022 | 9 | in | out
| HCHM20220816 | XO2-04 | CWeber | SPC | [5 empty] | 490.1 | 489.8 | 489.9 | 489.9 | 2764.9 | 2767 | 2022-09-20 17:15 | 09/21/2022 15:00 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.061212 | 2.1 | 0.075952 | 2.4 | 0.13716 | Excellent | 2022 | 9 | in | in
| MIR20220818 | XO2-56 | PScott | SPC | [5 empty] | 131.6 | 131.7 | 132.8 | 132.7 | 2763.4 | 2767 | 2022-09-20 17:15 | 09/21/2022 15:00 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.15198 | 3.6 | 0.13027 | 3.8 | 0.28225 | Excellent | 2022 | 9 | in | out
| SGA20220817 | XO2-01 | DColvin | SPC | [5 empty] | 1187 | 1188 | 1189 | 1194 | 2777 | 2767 | 2022-09-21 16:45 | 09/21/2022 15:00 | 5.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.33698 | -10 | -0.3601 | 14 | 0.69708 | Excellent | 2022 | 9 | in | out
| PDC20220817 | XO2-08 | DColvin | SPC | [5 empty] | 1190 | 1191 | 1185 | 1183 | 2765 | 2767 | 2022-09-21 17:30 | 09/21/2022 16:30 | 6 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.2521 | 2 | 0.072333 | 5 | 0.32443 | Excellent | 2022 | 9 | in | in
| TSL20220831 | XO2-69 | TSalman | SPC | [5 empty] | 2583 | 2583 | 2584 | 2590 | 2770.6 | 2767 | 2022-09-26 15:45 | 09/28/2022 10:45 | 50 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -6 | -0.23229 | -3.6 | -0.12994 | 9.6 | 0.36222 | Excellent | 2022 | 9 | in | out
| FCT20220831 | XO2-13 | TSalman | SPC | [5 empty] | 2577 | 2577 | 2574 | 2585 | 2768.9 | 2767 | 2022-09-26 16:15 | 09/28/2022 08:40 | 47.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -11 | -0.42685 | -1.9 | -0.068619 | 12.9 | 0.49547 | Excellent | 2022 | 9 | in | in
| FAL20220831 | XO2-43 | TSalman | SPC | [5 empty] | 2573 | 2569 | 2574 | 2572 | 2769.2 | 2767 | 2022-09-26 16:45 | 09/28/2022 09:50 | 48.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.07773 | -2.2 | -0.079445 | 4.2 | 0.15718 | Excellent | 2022 | 9 | in | in
| OSJ20220908 | XO2-26 | TSalman | SPC | [5 empty] | 2569 | 2568 | 2560 | 2556 | 2756.3 | 2767 | 2022-09-26 17:30 | 09/28/2022 08:45 | 46.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.11678 | 10.7 | 0.3882 | 13.7 | 0.50498 | Excellent | 2022 | 9 | in | out
| MOK20220831 | XO2-54 | TSalman | SPC | [5 empty] | 2552 | 2555 | 2557 | 2564 | 2772.7 | 2767 | 2022-09-26 18:00 | 09/26/2022 11:08 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.15674 | -5.7 | -0.20558 | 9.7 | 0.36232 | Excellent | 2022 | 9 | in | out
| STTD20220906 | ND-8 | EHuynh | SPC | [5 empty] | 186.9 | 187.2 | 187.1 | 187.2 | 2765.9 | 2767 | 2022-10-03 17:15 | 10/04/2022 14:15 | 28 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.10701 | 1.1 | 0.03977 | 1.3 | 0.14678 | Excellent | 2022 | 10 | in | in
| LIS20220906 | ND-5 | EHuynh | SPC | [5 empty] | 240.9 | 241 | 240 | 239.2 | 2758.4 | 2767 | 2022-10-03 18:15 | 10/04/2022 13:30 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.9 | 0.3736 | 8.6 | 0.31177 | 9.5 | 0.68537 | Excellent | 2022 | 10 | in | out
| RD2220220907 | ND-4 | EHuynh | SPC | [5 empty] | 850 | 850 | 848 | 850 | 2772.9 | 2767 | 2022-10-04 15:45 | 10/04/2022 14:50 | 6.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.23529 | -5.9 | -0.21277 | 7.9 | 0.44807 | Excellent | 2022 | 10 | in | out
| I8020220907 | ND-6 | EHuynh | SPC | [5 empty] | 674 | 674 | 669 | 669 | 2766.3 | 2767 | 2022-10-04 17:30 | 10/04/2022 15:20 | 4.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 0.7 | 0.025305 | 0.7 | 0.025305 | Excellent | 2022 | 10 | in | in
| GLC20220913 | XO2-09 | AMaguire | SPC | [5 empty] | 557 | 557 | 551 | 553 | 2754 | 2767 | 2022-10-05 16:15 | 10/06/2022 00:00 | 14.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.35907 | 13 | 0.47204 | 15 | 0.83111 | Excellent | 2022 | 10 | in | out
| ODM20220913 | XO2-49 | AMaguire | SPC | [5 empty] | 558 | 558 | 556 | 557 | 2772.4 | 2767 | 2022-10-05 16:45 | 10/06/2022 00:00 | 14.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.17921 | -5.4 | -0.19478 | 6.4 | 0.37399 | Excellent | 2022 | 10 | in | out
| ORM20220913 | XO2-16 | AMaguire | SPC | [5 empty] | 559 | 558 | 553 | 552 | 2753 | 2767 | 2022-10-05 17:15 | 10/06/2022 10:15 | 24 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 14 | 0.50854 | 14 | 0.50854 | Excellent | 2022 | 10 | in | out
| BET20220901 | XO2-61 | JFrantzich | SPC | [5 empty] | 970 | 970 | 966 | 969 | 2767 | 2767 | 2022-10-06 17:15 | 10/06/2022 10:17 | 0 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.30928 | 0 | 0 | 3 | 0.30928 | Excellent | 2022 | 10 | in | in
| BLP20220901 | XO2-53 | JFrantzich | SPC | [5 empty] | 1916 | 1915 | 1922 | 1912 | 6686 | 6668 | 2022-10-06 18:30 | 10/06/2022 14:30 | 3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 9 | 0.46973 | -18 | -0.26922 | 27 | 0.73895 | Excellent | 2022 | 10 | in | out
| ORI20220913 | XO2-55 | AMaguire | SPC | [5 empty] | 493.4 | 493.4 | 490.9 | 491.6 | 2763 | 2767 | 2022-10-10 16:30 | 10/10/2022 13:30 | 4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.7 | -0.14187 | 4 | 0.14477 | 4.7 | 0.28664 | Excellent | 2022 | 10 | in | out
| WCI20220913 | XO2-63 | AMaguire | SPC | [5 empty] | 493.4 | 493.3 | 488.7 | 490.7 | 2759.8 | 2767 | 2022-10-10 16:45 | 10/10/2022 13:45 | 4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2.1 | -0.42562 | 7.2 | 0.26089 | 9.3 | 0.68651 | Excellent | 2022 | 10 | in | out
| VCU20220913 | XO2-70 | AMaguire | SPC | [5 empty] | 493.3 | 493.8 | 488.9 | 491.1 | 2760.9 | 2767 | 2022-10-10 17:15 | 10/10/2022 14:00 | 3.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1.7 | -0.34462 | 6.1 | 0.22094 | 7.8 | 0.56556 | Excellent | 2022 | 10 | in | out
| TRN20220908 | XO2-35 | TSalman | SPC | [5 empty] | 514 | 514 | 509 | 513 | 2773.8 | 2767 | 2022-10-13 15:45 | 10/13/2022 17:20 | 8.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.77821 | -6.8 | -0.24515 | 10.8 | 1.0234 | Excellent | 2022 | 10 | in | out
| HLT20220908 | XO2-75 | TSalman | SPC | [5 empty] | 514 | 514 | 508 | 511 | 2764.3 | 2767 | 2022-10-13 16:30 | 10/13/2022 17:50 | 8.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.58366 | 2.7 | 0.097674 | 5.7 | 0.68133 | Excellent | 2022 | 10 | in | in
| WDC20220920 | XO2-30 | CWeber | SPC | [5 empty] | 598 | 599 | 596 | 599 | 2763.2 | 2767 | 2022-10-13 16:30 | 10/14/2022 09:45 | 24.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.33445 | 3.8 | 0.13752 | 5.8 | 0.47197 | Excellent | 2022 | 10 | in | out
| RSCC20220920 | XO2-68 | CWeber | SPC | [5 empty] | 600 | 603 | 599 | 602 | 2762 | 2767 | 2022-10-13 17:00 | 10/14/2022 09:45 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 5 | 0.18103 | 5 | 0.18103 | Excellent | 2022 | 10 | in | out
| ORQ20220823 | XO2-18 | TSalman | SPC | [5 empty] | 515 | 515 | 798 | 891 | 2959 | 2767 | 2022-10-13 17:15 | 10/13/2022 18:25 | 8.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -93 | -18.0583 | -192 | -6.4887 | 285 | 24.5469 | Poor | 2022 | 10 | out | out
| DRB20220920 | XO2-39 | CWeber | SPC | [5 empty] | 601 | 603 | 602 | 605 | 2771.4 | 2767 | 2022-10-13 17:30 | 10/14/2022 09:45 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16639 | -4.4 | -0.15876 | 5.4 | 0.32515 | Excellent | 2022 | 10 | in | out
| HOL20220823 | XO2-66 | TSalman | SPC | [5 empty] | 516 | 516 | 511 | 514 | 2770.4 | 2767 | 2022-10-13 17:45 | 10/13/2022 17:40 | 6.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.5814 | -3.4 | -0.12273 | 6.4 | 0.70412 | Excellent | 2022 | 10 | in | out
| HCHM20220920 | XO2-23 | CWeber | SPC | [5 empty] | 604 | 606 | 601 | 605 | 2766.3 | 2767 | 2022-10-13 18:00 | 10/14/2022 09:45 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.33113 | 0.7 | 0.025305 | 2.7 | 0.35643 | Excellent | 2022 | 10 | in | in
| OBI20220823 | XO2-02 | TSalman | SPC | [5 empty] | 830 | 831 | 816 | 830 | 2773 | 2767 | 2022-10-13 19:15 | 10/13/2022 18:04 | 5.8 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -13 | -1.5663 | -6 | -0.21637 | 19 | 1.7826 | Excellent | 2022 | 10 | in | out
| SJD20220907 | XO2-72 | BJones | SPC | [5 empty] | 168.3 | 169.7 | 169.2 | 169.2 | 2760.7 | 2767 | 2022-10-17 16:00 | 10/18/2022 12:40 | 27.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.4 | 0.83185 | 6.3 | 0.2282 | 7.7 | 1.0601 | Excellent | 2022 | 10 | in | out
| OH120220907 | XO2-15 | BJones | SPC | [5 empty] | 170.9 | 171 | 172 | 171.5 | 2767.6 | 2767 | 2022-10-17 16:30 | 10/18/2022 12:05 | 26.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.35108 | -0.6 | -0.021679 | 1.2 | 0.37276 | Excellent | 2022 | 10 | in | in
| ORX20220907 | XO2-19 | BJones | SPC | [5 empty] | 172.4 | 172.4 | 173.7 | 172.7 | 2759.3 | 2767 | 2022-10-17 17:15 | 10/18/2022 10:20 | 24.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.58005 | 7.7 | 0.27906 | 8.7 | 0.8591 | Excellent | 2022 | 10 | in | out
| GLE20220913 | XO2-03 | BJones | SPC | [5 empty] | 175.8 | 176 | 177.6 | 176.1 | 2760.3 | 2767 | 2022-10-17 17:45 | 10/18/2022 11:15 | 24.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.7 | 0.96701 | 6.7 | 0.24273 | 8.4 | 1.2097 | Excellent | 2022 | 10 | in | out
| TRB820220914 | XO2-21 | AMaguire | SPC | [5 empty] | 141.1 | 140.7 | 141.1 | 141 | 2767 | 2767 | 2022-10-18 15:00 | 10/18/2022 08:10 | 0.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.3 | -0.21262 | 0 | 0 | 0.3 | 0.21262 | Excellent | 2022 | 10 | in | in
| TRTT20220914 | XO2-25 | AMaguire | SPC | [5 empty] | 141.4 | 140.9 | 141.2 | 140.9 | 2763 | 2767 | 2022-10-18 16:15 | 10/19/2022 09:00 | 23.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.2 | -0.14144 | 4 | 0.14477 | 4.2 | 0.28621 | Excellent | 2022 | 10 | in | out
| LAB20220914 | XO2-41 | AMaguire | SPC | [5 empty] | 141.1 | 140.9 | 141.3 | 140.9 | 2765 | 2767 | 2022-10-18 17:15 | 10/19/2022 09:40 | 23.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.14174 | 2 | 0.072333 | 2.2 | 0.21408 | Excellent | 2022 | 10 | in | in
| FAR20220914 | XO2-65 | AMaguire | SPC | [5 empty] | 141.5 | 141 | 142.3 | 141.6 | 2768 | 2767 | 2022-10-18 18:15 | 10/19/2022 10:00 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.14134 | -1 | -0.036127 | 1.2 | 0.17747 | Excellent | 2022 | 10 | in | in
| MHO20220914 | XO2-59 | BJones | SPC | [5 empty] | 1005 | 1004 | 998 | 999 | 2756.3 | 2767 | 2022-10-19 16:00 | 10/20/2022 12:30 | 27.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.199 | 10.7 | 0.3882 | 12.7 | 0.58721 | Excellent | 2022 | 10 | in | out
| MRU20220914 | XO2-67 | BJones | SPC | [5 empty] | 1005 | 1004 | 995 | 991 | 2750.6 | 2767 | 2022-10-19 16:30 | 10/20/2022 12:55 | 27.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.29851 | 16.4 | 0.59623 | 19.4 | 0.89474 | Excellent | 2022 | 10 | in | out
| SGA20220921 | XO2-14 | DColvin | SPC | [5 empty] | 1172 | 1173 | 1169 | 1169 | 2772 | 2767 | 2022-10-19 16:30 | 10/20/2022 15:00 | 29.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.085324 | -5 | -0.18038 | 6 | 0.2657 | Excellent | 2022 | 10 | in | out
| PDC20220921 | XO2-74 | DColvin | SPC | [5 empty] | 1175 | 1174 | 1159 | 1168 | 2767 | 2767 | 2022-10-19 17:00 | 10/19/2022 15:00 | 5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -10 | -0.85106 | 0 | 0 | 10 | 0.85106 | Excellent | 2022 | 10 | in | in
| MRX20220914 | XO2-44 | BJones | SPC | [5 empty] | 1005 | 1005 | 997 | 995 | 2745.1 | 2767 | 2022-10-19 17:15 | 10/20/2022 13:45 | 27.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.199 | 21.9 | 0.79779 | 23.9 | 0.99679 | Excellent | 2022 | 10 | in | out
| TWA20220914 | XO2-40 | BJones | SPC | [5 empty] | 1005 | 1005 | 1005 | 1001 | 2771.2 | 2767 | 2022-10-19 18:00 | 10/20/2022 13:25 | 26.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4 | 0.39801 | -4.2 | -0.15156 | 8.2 | 0.54957 | Excellent | 2022 | 10 | in | out
| SXS20220920 | XO2-46 | PScott | SPC | [5 empty] | 140.6 | 140.8 | 141.2 | 141.2 | 2764.2 | 2767 | 2022-10-20 15:45 | 10/20/2022 13:30 | 4.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.2 | 0.14225 | 2.8 | 0.1013 | 3 | 0.24354 | Excellent | 2022 | 10 | in | in
| SOI20220920 | XO2-48 | PScott | SPC | [5 empty] | 141 | 141.4 | 141.4 | 141.1 | 2757.7 | 2767 | 2022-10-20 16:15 | 10/20/2022 13:30 | 4.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.7 | 0.49645 | 9.3 | 0.33724 | 10 | 0.83369 | Excellent | 2022 | 10 | in | out
| MIR20220920 | XO2-60 | PScott | SPC | [5 empty] | 142 | 142.1 | 142.1 | 142.3 | 2760.2 | 2767 | 2022-10-20 17:00 | 10/20/2022 13:30 | 3.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.070423 | 6.8 | 0.24636 | 6.9 | 0.31678 | Excellent | 2022 | 10 | in | out
| TSL20220926 | XO2-27 | TSalman | SPC | [5 empty] | 2436 | 2436 | 2437 | 2440 | 2773.1 | 2767 | 2022-10-25 15:00 | 10/26/2022 15:40 | 31.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.12315 | -6.1 | -0.21997 | 9.1 | 0.34312 | Excellent | 2022 | 10 | in | out
| FCT20220926 | XO2-11 | TSalman | SPC | [5 empty] | 2431 | 2430 | 2423 | 2430 | 2770.2 | 2767 | 2022-10-25 15:30 | 10/26/2022 15:20 | 30.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -8 | -0.32908 | -3.2 | -0.11552 | 11.2 | 0.4446 | Excellent | 2022 | 10 | in | out
| FAL20220926 | XO2-20 | TSalman | SPC | [5 empty] | 2425 | 2426 | 2431 | 2427 | 2768.3 | 2767 | 2022-10-25 16:00 | 10/26/2022 16:50 | 31.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5 | 0.20619 | -1.3 | -0.04696 | 6.3 | 0.25315 | Excellent | 2022 | 10 | in | in
| OSJ20220926 | XO2-34 | TSalman | SPC | [5 empty] | 604 | 606 | 618 | 606 | 2769.1 | 2767 | 2022-10-25 16:45 | 10/26/2022 17:20 | 31.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 14 | 2.3179 | -2.1 | -0.075837 | 16.1 | 2.3937 | Excellent | 2022 | 10 | in | in
| ORM20221005 | XO2-36 | AMaguire | SPC | [5 empty] | 989 | 990 | 989 | 987 | 2770.6 | 2767 | 2022-11-02 16:45 | 11/03/2022 09:15 | 23.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.30334 | -3.6 | -0.12994 | 6.6 | 0.43327 | Excellent | 2022 | 11 | in | out
| ODM20221005 | XO2-12 | AMaguire | SPC | [5 empty] | 988 | 989 | 986 | 989 | 2769.1 | 2767 | 2022-11-02 17:15 | 11/03/2022 08:50 | 22.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.20243 | -2.1 | -0.075837 | 4.1 | 0.27827 | Excellent | 2022 | 11 | in | in
| GLC20221005 | XO2-10 | AMaguire | SPC | [5 empty] | 980 | 986 | 974 | 985 | 2774.8 | 2767 | 2022-11-02 18:00 | 11/03/2022 09:45 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -5 | -0.5102 | -7.8 | -0.2811 | 12.8 | 0.79131 | Excellent | 2022 | 11 | in | out
| RD2220221004 | ND-5 | EHuynh | SPC | [5 empty] | 601 | 602 | 604 | 602 | 2782.5 | 2767 | 2022-11-03 15:00 | 11/04/2022 07:14 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE |  |  |  |  | 3 | 0.49917 | -15.5 | -0.55705 | 18.5 | 1.0562 | Excellent | 2022 | 11 | in | out
| I8020221004 | ND-8 | EHuynh | SPC | [5 empty] | 495.9 | 496 | 498.9 | 499.9 | 2797.3 | 2767 | 2022-11-03 16:30 | 11/03/2022 15:40 | 6.2 | NCRO | FALSE | FALSE | TRUE | TRUE |  |  |  |  | -0.9 | -0.18149 | -30.3 | -1.0832 | 31.2 | 1.2647 | Excellent | 2022 | 11 | in | out
| STTD20221003 | ND-3 | EHuynh | SPC | [5 empty] | 197.3 | 198.1 | 197.9 | 197.8 | 2766.9 | 2767 | 2022-11-03 17:45 | 11/04/2022 15:30 | 28.8 | NCRO | FALSE | FALSE | TRUE | TRUE |  |  |  |  | 0.9 | 0.45616 | 0.1 | 0.0036142 | 1 | 0.45977 | Excellent | 2022 | 11 | in | in
| LIS20221003 | ND-1 | EHuynh | SPC | [5 empty] | 324.2 | 324.6 | 325.7 | 324.9 | 2772.1 | 2767 | 2022-11-07 19:15 | 11/07/2022 11:22 | 0.1 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1.2 | 0.37014 | -5.1 | -0.18398 | 6.3 | 0.55412 | Excellent | 2022 | 11 | in | out
| VCU20221010 | XO2-49 | AMaguire | SPC | [5 empty] | 595 | 596 | 594 | 592 | 2765 | 2767 | 2022-11-08 18:30 | 11/09/2022 09:55 | 23.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.5042 | 2 | 0.072333 | 5 | 0.57653 | Excellent | 2022 | 11 | in | in
| WCI20221010 | XO2-16 | AMaguire | SPC | [5 empty] | 595 | 595 | 590 | 591 | 2766.1 | 2767 | 2022-11-08 19:00 | 11/09/2022 09:20 | 22.3 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16807 | 0.9 | 0.032537 | 1.9 | 0.2006 | Excellent | 2022 | 11 | in | in
| ORI20221010 | XO2-09 | AMaguire | SPC | [5 empty] | 595 | 595 | 593 | 594 | 2769 | 2767 | 2022-11-08 19:30 | 11/09/2022 09:40 | 22.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.16807 | -2 | -0.072228 | 3 | 0.2403 | Excellent | 2022 | 11 | in | in
| TRN20221013 | XO2-42 | TSalman | SPC | [5 empty] | 505 | 505 | 504 | 503 | 2764.9 | 2767 | 2022-11-09 18:15 | 11/10/2022 14:45 | 28.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.19802 | 2.1 | 0.075952 | 3.1 | 0.27397 | Excellent | 2022 | 11 | in | in
| HLT20221013 | XO2-71 | TSalman | SPC | [5 empty] | 506 | 505 | 506 | 506 | 2769.4 | 2767 | 2022-11-09 18:45 | 11/10/2022 16:15 | 29.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.19763 | -2.4 | -0.086661 | 3.4 | 0.28429 | Excellent | 2022 | 11 | in | in
| HOL20221013 | XO2-54 | TSalman | SPC | [5 empty] | 507 | 507 | 507 | 507 | 2760.6 | 2767 | 2022-11-09 19:45 | 11/10/2022 16:00 | 28.2 | NCRO | TRUE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 6.4 | 0.23183 | 6.4 | 0.23183 | Excellent | 2022 | 11 | in | out
| OBI20221013 | XO2-32 | TSalman | SPC | [5 empty] | 509 | 509 | 508 | 509 | 2766.3 | 2767 | 2022-11-09 20:15 | 11/10/2022 15:00 | 26.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.19646 | 0.7 | 0.025305 | 1.7 | 0.22177 | Excellent | 2022 | 11 | in | in
| MHO20221019 | XO2-19 | BJones | SPC | [5 empty] | 769 | 769 | 772 | 768 | 2755.2 | 2767 | 2022-11-15 17:15 | 11/16/2022 09:45 | 24.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 4 | 0.52016 | 11.8 | 0.42828 | 15.8 | 0.94844 | Excellent | 2022 | 11 | in | out
| MRU20221019 | XO2-03 | BJones | SPC | [5 empty] | 770 | 769 | 766 | 767 | 2759.8 | 2767 | 2022-11-15 17:45 | 11/16/2022 09:45 | 24 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.25974 | 7.2 | 0.26089 | 9.2 | 0.52063 | Excellent | 2022 | 11 | in | out
| MRX20221019 | XO2-15 | BJones | SPC | [5 empty] | 770 | 770 | 771 | 769 | 2753.9 | 2767 | 2022-11-15 18:30 | 11/16/2022 09:45 | 23.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 2 | 0.25974 | 13.1 | 0.47569 | 15.1 | 0.73543 | Excellent | 2022 | 11 | in | out
| TRB820221018 | XO2-18 | TSalman | SPC | [5 empty] | 73.4 | 73.5 | 73.6 | 73.2 | 2757.1 | 2767 | 2022-11-15 19:00 | 11/16/2022 09:00 | 22 | NCRO | FALSE | FALSE |  | TRUE |  |  |  |  | 0.5 | 0.6812 | 9.9 | 0.35907 | 10.4 | 1.0403 | Excellent | 2022 | 11 | in | out
| TWA20221019 | XO2-70 | BJones | SPC | [5 empty] | 775 | 770 | 770 | 769 | 2755.3 | 2767 | 2022-11-15 19:15 | 11/16/2022 09:45 | 22.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -4 | -0.51613 | 11.7 | 0.42464 | 15.7 | 0.94077 | Excellent | 2022 | 11 | in | out
| LAB20221018 | XO2-61 | TSalman | SPC | [5 empty] | 73 | 73.2 | 71.1 | 71.2 | 2761.3 | 2767 | 2022-11-15 20:30 | 11/16/2022 09:40 | 21.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.13699 | 5.7 | 0.20642 | 5.8 | 0.34341 | Excellent | 2022 | 11 | in | out
| FAR20221018 | XO2-66 | TSalman | SPC | [5 empty] | 179.1 | 179.4 | 178.9 | 178.6 | 2760.5 | 2767 | 2022-11-15 22:00 | 11/16/2022 10:35 | 20.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.6 | 0.33501 | 6.5 | 0.23546 | 7.1 | 0.57047 | Excellent | 2022 | 11 | in | out
| TRTT20221018 | XO2-58 | TSalman | SPC | [5 empty] | 179.5 | 179.6 | 178.9 | 179.1 | 2763.7 | 2767 | 2022-11-15 23:30 | 11/16/2022 10:00 | 18.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.05571 | 3.3 | 0.11941 | 3.4 | 0.17512 | Excellent | 2022 | 11 | in | out
| SGA20221019 | XO2-08 | DColvin | SPC | [5 empty] | 1387 | 1391 | 1387 | 1386 | 2751 | 2767 | 2022-11-16 18:45 | 11/17/2022 16:00 | 29.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 5 | 0.36049 | 16 | 0.58161 | 21 | 0.9421 | Excellent | 2022 | 11 | in | out
| PDC20221019 | XO2-23 | DColvin | SPC | [5 empty] | 1387 | 1388 | 1387 | 1395 | 2777 | 2767 | 2022-11-16 19:30 | 11/16/2022 16:00 | 4.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -7 | -0.50469 | -10 | -0.3601 | 17 | 0.86479 | Excellent | 2022 | 11 | in | out
| SOI20221020 | XO2-43 | PScott | SPC | [5 empty] | 158.4 | 159.2 | 158.4 | 158.7 | 2751.2 | 2767 | 2022-11-17 19:00 | 11/18/2022 12:25 | 25.4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.5 | 0.31566 | 15.8 | 0.57429 | 16.3 | 0.88995 | Excellent | 2022 | 11 | in | out
| MIR20221020 | XO2-56 | PScott | SPC | [5 empty] | 158.8 | 159.1 | 159.5 | 159.7 | 2755.9 | 2767 | 2022-11-17 19:45 | 11/18/2022 12:25 | 24.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.062972 | 11.1 | 0.40277 | 11.2 | 0.46574 | Excellent | 2022 | 11 | in | out
| HCHM20221013 | XO2-04 | PScott | SPC | [5 empty] | 701 | 702 | 696 | 698 | 2766.2 | 2767 | 2022-11-22 18:30 | 11/22/2022 16:00 | 5.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.14265 | 0.8 | 0.028921 | 1.8 | 0.17157 | Excellent | 2022 | 11 | in | in
| WDC20221013 | XO2-47 | PScott | SPC | [5 empty] | 702 | 702 | 699 | 700 | 2768.6 | 2767 | 2022-11-22 19:00 | 11/22/2022 16:00 | 5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.14245 | -1.6 | -0.057791 | 2.6 | 0.20024 | Excellent | 2022 | 11 | in | in
| DRB20221013 | XO2-50 | PScott | SPC | [5 empty] | 703 | 703 | 698 | 695 | 2765.7 | 2767 | 2022-11-22 19:30 | 11/22/2022 16:00 | 4.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 3 | 0.42674 | 1.3 | 0.047004 | 4.3 | 0.47375 | Excellent | 2022 | 11 | in | in
| RSCC20221013 | XO2-73 | PScott | SPC | [5 empty] | 704 | 704 | 702 | 702 | 2764.3 | 2767 | 2022-11-22 20:00 | 11/22/2022 16:00 | 4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 2.7 | 0.097674 | 2.7 | 0.097674 | Excellent | 2022 | 11 | in | in
| LIS20221107 | ND-6 | EHuynh | SPC | [5 empty] | 432.1 | 432.2 | 433.2 | 432.5 | 2778.5 | 2767 | 2022-11-28 22:15 | 11/29/2022 09:10 | 18.9 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.18514 | -11.5 | -0.41389 | 12.3 | 0.59903 | Excellent | 2022 | 11 | in | out
| SJD20221017 | XO2-01 | BJones | SPC | [5 empty] | 523 | 522 | 523 | 525 | 2750.1 | 2767 | 2022-12-02 17:30 | 12/06/2022 11:45 | 98.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.57361 | 16.9 | 0.61452 | 19.9 | 1.1881 | Excellent | 2022 | 12 | in | out
| OH120221017 | XO2-55 | BJones | SPC | [5 empty] | 522 | 522 | 523 | 523 | 2754.5 | 2767 | 2022-12-02 18:00 | 12/03/2022 10:30 | 24.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 12.5 | 0.4538 | 12.5 | 0.4538 | Excellent | 2022 | 12 | in | out
| ORX20221017 | XO2-53 | BJones | SPC | [5 empty] | 523 | 523 | 511 | 514 | 2798.4 | 2767 | 2022-12-02 18:30 | 12/06/2022 09:10 | 94.7 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.57361 | -31.4 | -1.1221 | 34.4 | 1.6957 | Excellent | 2022 | 12 | in | out
| GLE20221017 | XO2-63 | BJones | SPC | [5 empty] | 523 | 523 | 525 | 525 | 2759.2 | 2767 | 2022-12-02 19:00 | 12/06/2022 09:45 | 94.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0 | 0 | 7.8 | 0.28269 | 7.8 | 0.28269 | Excellent | 2022 | 12 | in | out
| MHO20221115 | XO2-48 | BJones | SPC | [5 empty] | 189.4 | 189.5 | 191.2 | 190.3 | 2756.1 | 2767 | 2022-12-07 17:15 | 12/08/2022 12:50 | 27.6 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.52798 | 10.9 | 0.39549 | 11.9 | 0.92347 | Excellent | 2022 | 12 | in | out
| MRU20221115 | XO2-09 | BJones | SPC | [5 empty] | 189.1 | 189 | 190.1 | 189.2 | 2756.4 | 2767 | 2022-12-07 18:15 | 12/08/2022 12:30 | 26.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.8 | 0.42306 | 10.6 | 0.38456 | 11.4 | 0.80762 | Excellent | 2022 | 12 | in | out
| MRX20221115 | XO2-49 | BJones | SPC | [5 empty] | 188.3 | 188.4 | 189.5 | 188.6 | 2760 | 2767 | 2022-12-07 19:00 | 12/08/2022 12:15 | 25.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.53107 | 7 | 0.25362 | 8 | 0.78469 | Excellent | 2022 | 12 | in | out
| TWA20221115 | XO2-16 | BJones | SPC | [5 empty] | 187.8 | 188 | 194.5 | 187.1 | 2762.1 | 2767 | 2022-12-07 19:30 | 12/08/2022 11:50 | 24.3 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 7.6 | 4.0469 | 4.9 | 0.1774 | 12.5 | 4.2243 | Good | 2022 | 12 | in | out
| GLC20221102 | XO2-12 | AMaguire | SPC | [5 empty] | 553 | 553 | 550 | 551 | 2759 | 2767 | 2022-12-13 18:15 | 12/14/2022 09:00 | 22.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18083 | 8 | 0.28996 | 9 | 0.47079 | Excellent | 2022 | 12 | in | out
| ORI20221108 | XO2-10 | AMaguire | SPC | [5 empty] | 554 | 553 | 551 | 553 | 2761 | 2767 | 2022-12-13 18:45 | 12/14/2022 09:00 | 22.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -3 | -0.54152 | 6 | 0.21731 | 9 | 0.75883 | Excellent | 2022 | 12 | in | out
| WCI20221108 | XO2-36 | AMaguire | SPC | [5 empty] | 555 | 554 | 552 | 552 | 2760 | 2767 | 2022-12-13 19:00 | 12/14/2022 09:00 | 22 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18018 | 7 | 0.25362 | 8 | 0.4338 | Excellent | 2022 | 12 | in | out
| VCU20221108 | XO2-44 | AMaguire | SPC | [5 empty] | 554 | 555 | 552 | 552 | 2763 | 2767 | 2022-12-13 19:30 | 12/14/2022 09:45 | 22.2 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 1 | 0.18051 | 4 | 0.14477 | 5 | 0.32528 | Excellent | 2022 | 12 | in | out
| ODM20221102 | XO2-60 | AMaguire | SPC | [5 empty] | 555 | 555 | 552 | 553 | 2757 | 2767 | 2022-12-13 20:15 | 12/14/2022 09:45 | 21.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -1 | -0.18018 | 10 | 0.36271 | 11 | 0.54289 | Excellent | 2022 | 12 | in | out
| ORM20221102 | XO2-67 | AMaguire | SPC | [5 empty] | 558 | 558 | 553 | 555 | 2753 | 2767 | 2022-12-13 20:45 | 12/14/2022 09:45 | 21 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -2 | -0.35842 | 14 | 0.50854 | 16 | 0.86696 | Excellent | 2022 | 12 | in | out
| SXS20221117 | XO2-03 | PScott | SPC | [5 empty] | 187.5 | 187.6 | 187.4 | 187.4 | 2767.1 | 2767 | 2022-12-20 18:30 | 12/20/2022 16:00 | 5.5 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | 0.1 | 0.053333 | -0.1 | -0.0036139 | 0.2 | 0.056947 | Excellent | 2022 | 12 | in | in
| SOI20221117 | XO2-15 | PScott | SPC | [5 empty] | 187.7 | 188 | 188.1 | 188.5 | 2769.6 | 2767 | 2022-12-20 19:15 | 12/20/2022 16:00 | 4.8 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.053277 | -2.6 | -0.093876 | 2.7 | 0.14715 | Excellent | 2022 | 12 | in | in
| MIR20221117 | XO2-70 | PScott | SPC | [5 empty] | 188.2 | 188.2 | 188.9 | 189 | 2763.1 | 2767 | 2022-12-20 20:00 | 12/20/2022 16:00 | 4 | NCRO | FALSE | FALSE | TRUE | TRUE | TRUE |  |  |  | -0.1 | -0.053135 | 3.9 | 0.14115 | 4 | 0.19428 | Excellent | 2022 | 12 | in | out